Nitrogen in baron nitride ceramics

Power mode:					On Off On mode

Sample drop:					Manually

Out-gassing (Power / time)			5.0 kW for 60s

Purge time:						15s

Stabiliser time:					30s

Integration delay:					2s

Analysis power / time:				50 kW for 60s

Min. analysis time:				100s

Accelerator / Flux:					None

Calibration:						Pure chemicals (BN)

Oxygen in magnesium

Power mode:					On Off On mode

Sample drop:					Automatic

Out-gassing (Power / time)			4,5 kW for 45s

Purge time:						10s

Stabiliser time:					20s

Integration delay:					10s

Analysis power / time:				2,8 kW for 30s

Min. analysis time:				50s

Accelerator / Flux:					1g Sn

Calibration:						Standard copper.

Furnace cleaning after each analysis

Oxygen in steel

Power mode:					On Off On mode

Sample drop:					Automatic

Out-gassing (Power / time)			5.0 kW for 30s - 45s

Purge time:						10s

Stabiliser time:					20s

Integration delay:					10s

Analysis power / time:				4,0 kW for 30s

Min. analysis time:				50s

Accelerator / Flux:					None

Calibration:						Standard steel O+N in similar range.

Oxygen and nitrogen in titanium

Power mode:					On Off On mode

Sample drop:					Automatic

Out-gassing (Power / time)			5,8 kW for 45s

Purge time:						15s

Stabiliser time:					30s

Integration delay:					10s

Analysis power / time:				5,0 kW for 45s

Min. analysis time:				50s

Accelerator / Flux:					(250mg) sample placed in a 1g nickel basket

Calibration:						Standard titanium 100 ppm. N / 1000 ppm O.

High crucibles required.

Oxygen nitrogen in zirconium

Power mode:					On Off On mode

Sample drop:					Automatic

Out-gassing (Power / time)			5.5 kW for 30s - 45s

Purge time:						10s

Stabiliser time:					30s

Integration delay:					10s

Analysis power / time:				4,5 kW for 45s

Min. analysis time:				50s

Accelerator / Flux:					(250mg) sample placed in a 1g nickel basket

Calibration:						Standard titanium or zirconium.

High crucibles required.

Oxygen and nitrogen in ceramics

Power mode:					On Off On mode

Sample drop:					Manually

Out-gassing (Power / time)			5,0 kW for 45s

Purge time:						10s

Stabiliser time:					30s

Integration delay:					2s

Analysis power / time:				3,0 kW for 60s

Min. analysis time:				100s

Accelerator / Flux:					None

Calibration:						Pure chemicals in similar range.

10 – 100mg sample weight depending on

oxygen concentration.