To replace the manifold part # 11822 :(picture 1 below)

You need a wrench and metric Allen keys (they were in the small toolbox that came with the analyzer). Shut down the instrument, open the right side, use a big Allen key to disconnect the manifold from the chassis, use a wrench to disconnect the valves. Make sure you mark them and put them back in the same position.

Shut down the compressed air and remove the line from the back (to remove pressure). Disconnect all the tubes and connect to the new manifold.

Then reconnect the electrical part and reconnect compressed air.  Put back the valves in the same order. Reconnect the manifold to the chassis using the big Allen key.