Tutorial for removing the soot from the Belec Compact port

You need a rubber stopper and a small flat screwdriver.

1. First, shut down the spark voltage. There is a big black switch on the back of the instrument, out in position 0.

2. Place the rubber stopper on the sparking stand

3. Increase the flow using the adjustment placed on the big connector, rotate counter-clockwise to release the flow

4. Press the rubber stopper on the sparking stand and with the right hand block the flow at the exhaust

5. After a few seconds, release the exhaust and the soot will be pushed out

6. Repeat two or three times until there is no more soot coming

7. Switch back the spark voltage to 1 (reversing step 1 above)

8. Reduce the stand-by flow to 0.5-1 by turning clockwise the adjustment (reversing step 3 above)