Flow problems at CS-800 and CS-2000

Close the piston (F2)

Log in a power user by clicking on Start Uni, or from Programs, Uni, Start Uni

Click on "Password", enter "etna"

Enter the password

Open Device State Window

When you're done, just click on the two icons Configuration and Device State and the two windows will close. Then start using the instrument.

# Text boxes

Device state

Configuration

The flow in litres/hour, needs to be 180 l/h

Select this number and then adjust the voltage up and down until you see the flow in the right window, reach  180 l/h, +-5