NOTE :

Servicing should only be done by a person with profession, who has mastered the servicing and maintenance of this device, as well as having further qualifications, especially in the field of electronics and physics.

We ask  the users of this Service Manual to inform us about any possible mistakes.

We would also appreciate any suggestions for supplements and improvements to this Service Manual.

CAUTION !

Before opening this device, always unplug the main power !

## Index

Service  Manual  CS-2000

0.1	Overview of  the CS-2000 modifications..............................................................0.1-5

1	  Faults

1.1	IR-cell test display  ++++  see  2.  3 	1.1-1

1.2	Gas flow too high or too low	1.2-1

1.3	The gas flow rate fluctuates	1.3-1

1.4	No combustion	1.4-1

1.5	Bad combustion	1.5-1

1.6	Electronic drivers malfunctioning	1.6-1

1.7	The induction furnace does not close	1.7-1

1.8	The analysis takes too long	1.8-1

1.9	Development of dust when closing the furnace 	1.9-1

1.10	Erratic results ........	1.10-1

1.11	The results are always 0.0000% or 00.0 ppm	1.11-1

1.12	Beeper responding 	1.12-1

1.13	No gas flow 	1.13-1

1.14	The LED for the infrared cell temperature regulation does not flash	1.14-1

1.15	The fuse burns out at the beginning of combustion	1.15-1

1.16	The fuse blows at the end of the analysis	1.16-1

1.17	Large collection of dust when opening the furnace.	1.17-1

1.18	No or low oxygen pressure	1.18-1

1.19	Saved data are lost 	1.19-1

1.20	Noise or drift of the infrared baseline	1.20-1

1.21	The life time of the combustion tube is too short 	1.21-1

1.22	The word ELTRA appears on the display	1.22-1

1.23	The oxygen pressure falls when the furnace opens	1.23-1

1.24	Combustion without  starting the analyses	1.24-1

1.25	The crucible cracks or melts	1.25-1

1.26	The analysis doesn´t start 	1.26-1

  - The oxygen pressure goes up very slowly 	1.27-1

1.28	Pump capacity vs maximum	1.28-1

1.29	Resistance furnace takes too long to warm 	1.29-1

1.30	Negative drift after the start 	1.30-1

1.31	The resistance furnace is completely cold or too hot 	1.31-1

1.32	Smoke in the resistance furnace 	1.32-1

1.33	Resistance furnace doesn’t reach full temperature 	1.33-1

1.34	Unstable resistance furnace temperature 	1.34-1

1.35	Dust or soot from the resistance furnace 	1.35-1

1.36	Vibration in induction furnace mode 	1.36-1

1.37	The analysis cycle stops only 25 seconds after start 	1.37-1

1.38	The chopper motor doesn’t rotate 	1.38-1

1.39	High carbon results are erratic 	1.39-1

2	  Adjustments, tests and function explanations

2.1	Gas flow controller adjustment and jumper settings	2.1-1

2.2	Pressure switch	2.2-1

2.3	Infrared zero-baseline, control and adjustment 	2.3-1

2.4	IR-source voltage setting	2.4-1

2.5	Temperature regulation of the IR-cell	2.5-1

2.6	Pneumatics	2.6-1

2.7	Pressure regulators	2.7-1

2.8	Closing cone, adjustment	2.8-1

2.9	Linearisation	2.9-1

2.10	Programming the balance	2.10-1

2.11	 	 2.11-1

2.12	Induction generator control	2.12-1

2.13	Diagnostic printouts 	2.13-1

2.14	Setting check up	2.14-1

2.15	Thermocouple testing.	2.15-1

2.16	Message "pump vs maximum" 	2.16-1

3	  Service

3.1	IR-paths, cleaning and replacing	3.1-1

3.2	IR-source	3.2-1

3.3	Infrared electronics	3.3-1

3.4	Infrared power supply	3.4-1

3.5	Chopper  ++++  see  3. 2 	3.5-1

3.6	Infrared temperature regulation	3.6-1

3.7	Leak checking	3.7-1

3.8	Replacing EPROMs	3.8-1

3.9	Clearing memory	3.9-1

3.10	Electronic reset 	3.10-1

3.11	Coding IR-ranges	 3.11-1

3.12	Adding new IR-ranges	3.12-1

3.13	Cleaning the solenoid valves	3.13-1

3.14	Replacing the flow sensor	3.14-1

3.15	Installing an AC-power stabiliser	3.15-1

3.16	Additional safety features for induction furnace pneumatics (optional)	3.16-1

3.17	Replacing the gas flow pump..	3.17-1

3.18	Replacing the heating elements	3.18-1

3.19	Replacing the combustion tube	3.19-1

3.20	Installing a halogen trap 	3.20-1

3.21	Gas saving mode 	3.21-1

3.22	Gas saving also for resistance furnace .	3.22-1

3.23	Upgrade for copper analysis without accelerators  	3.23-1

3.24	Upgrading of analysers with electronic rack for using the new

Windows software	3.24-1

4	  Miscellaneous

4.1	Ordering numbers	4.1-1

4.2	Wiring and gas flow diagrams	4.2-1

4.3	Starting through external key	.4.3-1

4.4	External report of analysis condition	4.3-1

4.5	Spare parts kits.	4.5-1

4.6	Maintenance CS-2000.	4.6-1

0.1   Overview of the CS-2000 modifications:

Up to S/N 0500xxxxxx

The furnaces have heating elements with thick ends ( 14 mm diameter ).

From S/N 0500xxxxxx up to 0709xxxxxx

The furnaces have heating elements with only one uniform diameter of 8 mm.

From S/N 0710xxxxxx

The furnaces have heating elements with thick ends ( 14 mm diameter ).

CAUTION:

The above serial numbers are the serial numbers of the furnaces, not of the analyzers.

From S/N 1501xxxxxx

The electronic unit is replaced by the board UNI 1.3 and the electronic rack is removed. The corresponding wiring diagram is 12061 and 24811.

The modification took place in the second half of 2004.

# 1 FAULTS

## 1.1  IR-cell test display:  ++++  see 2. 3

also after five minutes of oxygen flow:

If the range is higher than 	10 V, ++++

then the problems are as follows:

1). Too much contamination of

the IR-paths:

In case the analyser has more

than one cell, and if not all of

them show  ++++, then a

contamination is very likely.

2). The IR-Sources:

The sources are bent or

de-adjusted through

transport. See Remark:

3). Chopper:

The chopper is not running.

This is obvious when the paths

are clean and the IR-sources

are in place.

Remark:

In most cases the problem is

in the IR-sources or in the

power supply. During a rough

transportation the IR-sources

may be bent causing the

problem. See 3.2. We

experienced that it is very

popular for the new IR-detector

and preamplifiers. We can only help if you send

us a copy of the flow diagram of this chapter, after drawing a line along the path you followed according to the condition you found.

Asking for detectors, preamplifiers and boards without following this diagram is useless.

## 1.2  Gas flow too high or too low:

Valid for induction furnace operation:

4	Panel meter indicating the control of the gas regulation valve

5	Gas flow meter 300 l/h

6	Oxygen pressure gauge

7	Compressed air pressure gauge

## 1.3  The gas flow rate fluctuates:

The flow rate is electronically regulated, therefore there should be no significant variation from the set value.

In case of a fault however, the automatic regulation of the flow rate is no longer possible and the gas flow fluctuates. This can be caused, for example, by an excessive obstruction of the gas flow system or by a faulty component that needs to be replaced, see 1. 2 and

2. 1

Most of the time, the problem lies in the saturation of the dust trap or of the chemicals.

In case the flow becomes lower and the µA-meter goes higher during combustion, it is obvious that the flow regulator is working. Therefore the problem must be a restricted oxygen supply. When the accelerators start burning they take a lot of oxygen, so that the pressure in the furnace goes down, and as a result for this the flow also goes down.

The restriction can be in:

- the hole in the ceramic shield

- the furnace outlet

- the connection from the furnace outlet to the metal filter housing

- the inlet of the metal filter housing

- the metal dust filter

- the anhydrone

- the paper filter

- the filter cartridge

Please note that the restrictions can also be outside the analyser, for example a faulty pressure regulator on the oxygen bottle, a too low adjustment of the pressure on the regulator, not properly opened outlet tap of the regulator, a too long and too thin tube, or this tube is somewhere bent or pressed or, or, or.

If the problem is only a long thin tube, the problem can eventually be solved by increasing the pressure on the pressure regulator of the bottle.

As a hint, when you open the furnace the pressure on the analyser’s oxygen gauge should go from 1.5 bar down to 1.4 bar but not lower. If it goes much lower, it means that the incoming oxygen is restricted on the way from the bottle to the analyser.

## 1.4  No combustion:

Induction furnace operation

AFTER  PRESSING < START >

"ANALYSIS" AND "GEN" APPEARS, BUT COMBUSTION DOESN’T TAKE PLACE.

- The command from the electronic unit to start the generator, is not received:

Measurements on TP9:

- A signal of about 0.5 V should turn the generator on,

- A signal of about 12 V should turn the generator off.

If 12 V are measured, despite "GEN" appearing on the display, then the electronic is

defective, see 1. 6; or the connection from the HF 42 circuit board, pin 7 is loose,      see 4.2.

When the command for

the generator is received

( 0.5 V DC ), the needle pulse must

appear on the TP13 test point. If

there is no oscilloscope available

you can test this with a digital

voltmeter.

The displayed value depends on

what type of digital voltmeter you

have. There must be a voltage

increase on TP13, once the word "GEN"

has appeared on the display, or once

0.5 V are present on TP9. When no voltage increase occurs on TP13, despite having

0.5 V on TP9; the HF 42 circuit board is defective.

CAUTION! After replacing the HF 42 board, make sure that the 50 pin plug is

properly inserted. It takes a lot of force to push in 50 pins at once.

50 pin plugs that haven’t been properly inserted are a frequent problem.

Tighten the screws of the plug properly, to ensure a proper connection to the board.

Checking the generator tube:

- Ensure that the power switch is at setting 2

- Check if the filament of the generator tube glows:

This can be easily checked:

through the air vent holes of the furnace,

the reflected light of the glowing filaments

( heater of the cathode ) should be visible.

If the light from the filament is not visible, then remove the

cover from the generator section, in order to see the tube.

WARNING !  HIGH  VOLTAGE !

If the filament of the generator tube doesn’t glow:

- Check the voltage for the heater

on the transformer, there should be 10 V AC

on the secondary, and 220 V AC primary

voltage.

- Check whether the filament is faulty,

at the filament connections, only 1 Ohm

must be measured. A very high or

unmeasured resistance means that the

filament is burnt through and the tube needs

to be replaced.

CAUTION!

- Set the power switch to pos. 1 before removing

the tube, although it would be best

to completely turn off the device and unplug the power cord.

Checking the primary voltage on the high voltage transformer:

- On the primary side of  the high voltage transformer, you should be able to measure

210 V AC; if not, the triac is defective. ( To locate the triac, see 4.1-3 )

WARNING !  HIGH  VOLTAGE !

High voltage transformer:  windings check up:

- In case the transformer has 210 V AC on the primary coil and there is still no

combustion, check the windings of the transformer as follows:

- Switch off the analyser and disconnect the main power plug.

- Measure the resistance of the primary coil.

It must be extremely low, about 0.8 Ohms between the screws labelled 0 and 270.

- Measure the resistance of the secondary coil between the two screws on the ceramic

insulator where the wires of the rectifier are connected. The resistance is approximately 322 Ohms.

- The rectifier ( 1 ) could also be faulty!

CHECKING THE HIGH

FREQUENCY OSCILLATOR:

Warning  !  HIGH  VOLTAGE !

Never try to measure the voltage here !

As there is high voltage ( 5000 volts ) and

high frequency ( 19.2 MHz ) present, a

measurement would be very difficult and,

above all, very dangerous. Before the

generator cover plate is unscrewed, the

device should be turned off or, better, the

power plug pulled out.

- You should check the connections in the

oscillating circuit to ensure all screws are

turned tightly, especially all the screw caps

on the combustion coil.

- The electrical connections and resistors

should be checked with an ohmmeter.

- If the resistor 2.7 KΩ / 100 W ( 3 ) is

open circuit, then there is no oscillation.

The parts e.g. tube and rectifier should be

replaced one by one, see also 1. 15

Resistance furnace operation:

- The furnace temperature is too low.

See Operation manual 3. 1

- The temperature set point is set correctly, but the actual furnace temperature is incorrect. See 1.31 and 1.33

## 1.5  Bad combustion:

Induction furnace operation:

Insufficient flow: see 1. 3

- The bypass valve ( 28 ) is either falsely activated during combustion, or it is sticking.   In this case, the oxygen flows through the bypass valve into the IR-cell and not through the furnace. The furnace remains nonetheless under oxygen pressure.

- Measure the voltage to the bypass valve ( 28 )

24 Volts DC while the furnace is open is normal.

Zero volts when the furnace is closed is normal.

The measurements can also be made with a digital voltmeter, as well as with a glow light probe that shows 24 V / 20 mA.

Other false controls from the electronics :

Clear the memory, see 3. 9

Clean the valves, see 3.13

Generator tube is too old:

This normally does not occur suddenly, it is rather a long process.

- Replace the tube

Replace inadequate contact points in the oscillating circuit, see 1. 4

Replace worn power controls, see 1. 4

Resistance furnace operation:

- The furnace temperature is too low, see 1.31 and the Operation Manual 3. 1 and 3.4

- The sample contains a substance that doesn’t burn easily.

Some materials contain sulphur compounds wich are very difficult to burn, i.e. they need higher temperature than 1550°C to properly burn in a reasonable time.

An improvement can be achieved by adding accelerators, such as iron phosphate to coal samples that contain pyrites, or vanadium pentoxide to cement samples.

For materials which require such high temperatures to burn, we recommend induction furnace analysers.

No oxygen flow, see 1.32

## 1.6  Electronic- drivers malfunctioning:

IC a, b, c and d are electronic driver

integrated circuits, found on the mother

board inside the electronic unit.

Their Type No. is ULN 2803 or

TD 62083.

These IC’s act as a power switch, in

order to energise solenoid valves inside

the analyser, see 2. 6, 3. 7 and 4. 2

or to send control signals, see 1. 2

1. 4 and 1.24

Operation range of IC’s  a, b, c, d.

Depending on the fault location, the following driver IC’s need to be replaced:

IC a: 	Outflow valve V1, see 3. 7 and 4. 2

IC b: 	Shuttle valve V2, see 3. 7 and 4. 2 ; pneumatic valve for furnace, see 2. 6

IC c: 	Purging valve V3, see 3. 7and 4. 2

IC d: 	By-pass valve V4, see 4. 2

Generator „start“ signal, see 1. 4 and 1.24

## 1.7  The induction furnace does not close:

- Little or no compressed air pressure, the furnace cannot close.

This pressure is shown on gauge ( 11 ), on the front-side of the analyser.

It must be between 4 to 6 bar ( 60 to 90 psi ).

- The screws ( 10 ) and ( 3 ) must be turned counter-clockwise, as much as possible.

- Adjust regulator ( PR4 ), so that the furnace floor can rise completely, without shaking; while a person still can stop the moving furnace floor with only one hand.

This is a safety precaution.

Repeat above instruction, until the adjustment is completed. Now, the upward speed is fast enough, yet safe for the operator. The pressure regulator ( PR4 ) is found on the inside, on the right-side section of the analyser.

- The pneumatic valve ( V7 ) is malfunctioning.

- The magnetic coil of the pneumatic valve ( V7 ) receives no current.

The furnace floor goes up, when valve ( V7 ) receives 24 V DC.

The 24 VDC comes from the electronic unit; if not, replace the electronic drivers,

see 1. 6

## 1.8  The analysis takes too long:

The analysis length lasts a maximum of up to 100 seconds, by virtue of a maximum time limit which the operator can select.

The minimum analysis time is automatically defined by the analyser and corresponds to half the maximum time.

Once the minimum time has expired, the analysis will stop automatically, if by that time no combustion has taken place.

If however the required combustion does take place, a signal from the infrared cell  is then produced, causing the analysis time to be automatically lengthened.

The analysis is ended once the infrared signal has descended to the comparator        threshold level, which is nearly zero volts.

The comparator threshold level is an artificial zero level.

When, however, as a result of unwanted hydrocarbons in the carrier gas, there is too much CO2, this threshold level won’t be reached.

In this case the analysis will stop only once the selected maximum time has expired.

The purpose of the maximum time is to protect the furnace from overheating.

It is important to ensure that the analysis is completed before the maximum time has expired.

This threshold level is not the same for every analysis, it is reset automatically during each analysis.  The higher the maximum peaks, the higher the threshold level.

It can therefore happen that with higher peaks the analysis ends sooner than with lower peaks.

As long as the unwanted hydrocarbon concentration in the oxygen is not extremely high, the problem could only exist in the low carbon channel ( if present ).

You can confirm the source of this problem by deactivating the low carbon channel;

< Menu >  < 5 >  < 1 > .

If the analysis then ends before max. time, you can be fairly sure that the problem is with hydrocarbons.

In case that there is neither an oxygen purification furnace, nor some clean oxygen available, you can by-pass this problem by setting the threshold level higher. Press the keys;

< Menu >  < 9 >  < 1 >  < k >  < XXX >  < ENTER >.

< k >  can be the number 1, 2 or 3 for selecting the required IR- detector.

< XXX >  is a three digit number for the set point of the threshold level, for example 250.

If the problem remains, this number can be increased stepwise until the problem is solved. The threshold level can only be set in the low carbon channel.

Additional reasons for excessive analysis duration could be:

- Dirty dust trap, see Operation Manual 6. 4

- Saturated chemicals, see Operation Manual 7. 2

- Cotton, glass or quartz wool is pressed too tightly, see Operation Manual 7. 2

- The oxygen is not pure, or the catalyst furnace is faulty or is cold, or the copper oxide is used up, see Operation Manual 7. 2

- The temperature of the infrared cell is not yet stable. Either it is not yet operational for at least 1 to 2 hours, or the oxygen flow has not yet been running for 10 to 15 minutes ( power switch on setting 2 ), see Operation Manual 3. 1  and 3. 2

- The thermostatic control of the infrared cell is defective, see 1.14  and 2. 5

- Excessive combustion temperature, causing the crucible to emit CO2.

One of the IR-sources is unstable:

- Small resistance changes of one or more IR-sources can cause the

current to vary, thus affecting the intensity of the IR-radiation.

This problem will effect all the channels, due to the fact that the IR-sources are all

connected in series.

In order to establish whether this problem is present or not, an analysis cycle should be run without a sample or a crucible. The baseline of the IR-cells can then be monitored on the PC screen, see Operation Manual 5. 8 PLOTCS

The baselines should be quite stable and should run between 0.100 V and – 0.100 V

It would be very difficult, in case of a problem, to establish which IR-source is faulty and it is therefore easier to replace them all.

- A window of an IR-path is not airtight. The measuring gases pass through the leakage, either into the chopper housing or to the IR-detector.

Perform a leakage test of the IR-paths. Replace the faulty IR-path. See 3. 1

- The rectifier of the IR-source supply is faulty. The unstable voltage is too low, so that the voltage at the IR-sources cannot be regulated. see 2. 4

Check the base line on the PC screen, by starting analysis without sample. Zoom to get the scale of 10 mV steps. See Operation Manual 5. 8

The base line should remain within 10 mV, 50 seconds after START.

If the drift is larger and always positive, then check the IR-source voltage. See 2. 4

- The infrared power supply is defective, so that the voltage stabilisation does not function.

It is possible that the current to the IR-source is unstable, due to a faulty rectifier ( 4 ) causing only a half-wave to pass through. This can be checked rather easily, by measuring the DC-voltage across the outer legs of the rectifier ( DC-output ).

It should be at least 4 V DC higher than the supply voltage for the IR-source, measured between TP2 and TP1 ( GND ).

Voltage across DC output of the rectifier:

Voltage between TP1 ( GND ) and 		At least    6 V DC	for 1 channel device

capacitor C:		At least 7,5 V DC	for 2 channel

( not stabilized voltage before stabilizer )	At least 9,5 V DC	for 3 channel device

At least  11 V DC	for 4 channel device

IR-source voltage, is to

be measured between TP1

and TP2 on IR 45 board.			Transformer connections:

( stabilized voltage of IR – sources )

Blue cables

1,75 V DC   for  1 channel device		1  Soldering location for 1 channel device

3,5   V DC   for  2 channel device		2  Soldering location for 2 channel device

5,25 V DC   for  3 channel device		3  Soldering location for 3 channel device

7      V DC   for  4 channel device		4  Soldering location for 4 channel device

( 1,75 Volts DC for each IR source )	8  Always connected ( common )

- RF-interference on the IR-cell, or on the electronic units:

If the generator's high frequency disrupts the electronic units or the infrared cell, then

the base line on the PC monitor appears jerky.

CAUTION !  HIGH  VOLTAGE !

The causes:

- The screws ( 1 ) are not tightened properly, or not all 22 of them are screwed in.

- The acorn nuts ( 2 ) are not tightened properly or not all 14 are in place.

- The cover ( 3 ) does not seal tightly with the plates ( 4 ) and ( 5 ) or with

the front panel ( 6 )

- The capacitor ( 10 ) is out of order

- The capacitor ( 9 ) is out of order

One or both filters ( 7 ) are not OK:

- The filter ( 11 ) can also be damaged in this way.

When the ground strip ( 12 ) is not connected after removal or replacement of

the oscillator, the filters ( 7 ) will get defect.

The power voltage sinks below 180 VAC:

- Even a momentary drop in the voltage is enough to cause problems.

Use a voltage stabiliser. See 3.15

## 1.9  Development of dust when closing the furnace:

- The sample material is causing dust:

This will be caused mainly through tin analysis Tin causes a light grey dust. The same problem appears when tin is used as accelerator. Therefore using tin as accelerator should be avoided. Also the tungsten – tin mixture shouldn’t be used. Almost all different materials can be analysed without the need of tin accelerators.

See Operation Manual 3. 3 “Applications”

Usage of other materials:

Cast-iron causes brown dust during the analysis.

When the dust is yellow, too much tungsten is used as accelerator. If the dust is brown, too much Iron accelerator is used.

See Operation Manual 3. 3 “Applications”

## 1.10   Erratic results:

There are a lot of different reasons for this:

- Saturated chemicals or clogged dust filter,

see Operation Manual 6.2 and 6.3

- Heterogeneous samples

- Leakage, see 3. 7

- Flow problems, see 1. 2 , 1.3 and 1.18

- The electronic noise of the infrared signal, see 1.20

This can happen quite often that the combustion start is delayed and the analysis time is too long, see 1. 8

You will see that the combustion start is delayed, when the word "GEN" appears later than the word "ANALYSIS".

- Incomplete or bad combustion, see 1. 5

- Too high blank value for the crucible and tungsten. This will be especially noticeable,

mostly though, in the low carbon channel.

- Widely different weights of the sample and the accelerator ( tungsten ).

- Wrong data in the memory of the electronics: Clear memory, see 3. 9

- The base line of the detector is unstable, due through insufficient main voltage

( less than 200 V AC ).

Connect a stabiliser, see 3.15

- Bad combustion, see 1. 5 and also see Operation Manual 3. 4

- Wrong gas flow see, 1. 2 and 1. 3

Only the sulphur results are not repeatable:

When the Sulphur results are not stable, but the carbon results are good, there

are two possible reasons:

- The magnesium perchlorate ( anhydrone ) is saturated.

- Replace the complete contents of the glass tube after the corresponding furnace.

There is sulphuric acid in the combustion tube of the resistance furnace. Replace the

combustion tube, see 3. 19

- If the combustion tube ( 4 ) was in use for a very long time and many samples

with high Sulphur were analysed, then sulphuric acid may build up in the rear-side

of the combustion tube.

When the heated tube is removed together with the connector ( 5 ),

a brown fluid can be seen in the ceramic tube.

During the analysis, some SO2 is trapped in this fluid, causing non

repeatable results. It is recommended to clean the furnace outlet tube.

## 1.11   The results are always 0.0000% or 00.0 ppm:

## 1.12   Beeper responding:

The built-in beeper of the analyser responds due to the following reasons:

- When trying to analyse without entering the sample weight ( when the analyser operates without the PC )

- No oxygen input

- 24 V supply for the solenoid valves faulty

- No balance attached ( when analyser operating without PC )

At request, the beeper can be made to respond every time when an analysis is

completed.

## 1.13   No gas flow:

Resistance furnace operation

- For induction operation, see 1. 2 and 2. 1.

## 1.14   The LED for the IR-cell temperature regulation does not flash:

- Check that the cable is well connected to the plug, and that there are no short circuits between the contacts in the 13 pin plug.

- Inspect the spring contact ( 2 ) on the connector ( 1 )

- One of the contact threads ( 2 ) could have failed, or one of the connections ( 3 )

could be disconnected.

A look inside the plug ( 4 ) is recommended.

## 1.15   The fuse burns out at the beginning of combustion:

- If the fuses ( 11 ) or ( 12 ) blow, then unplug the main power and replace them.

- Disconnect the cables ( 1 ) from the primary side of the transformer ( 2 ) and

connect the cables ( 1 ) to a 220 V light bulb ( 3 ) which is inserted to a socket.

An ordinary desk lamp can serve for this purpose. In case only 110 V light bulbs

are available, connect two of them in series.

BE  CAREFUL  that  no short  circuit  occurs !

After the cables ( 1 ) have been connected, plug the analyser to the main power again.

- Set the power switch to pos. 2

- Enter a weight manually, for example, 500 mg

- Press < START >

- When "GEN" appears on the display, the light bulb should glow; this does not happen suddenly, but gradually.

- Within a few seconds, the light bulb will change from dark to fully bright.

WARNING !  HIGH VOLTAGE!

SEE DRAWING NEXT PAGE:

- After pressing < START >, the message "ANALYSIS" appears, but "GEN" could appear later. The reason is, that after restarting the device and the oxygen flow, the infrared cell base-line could drift.

For the same reason, the analysis may stop longer than usual.

- You should wait until "GEN" appears.

- Quit by pressing < STOP >. If the light bulb turns dim, the triac ( 6 ) is OK.

- Unplug the main power and reconnect the power cables ( 1 ) to the transformer ( 2 )

WARNING ! HIGH  VOLTAGE !

When reconnecting the power cables ( 1 ), do not let these wires approach the connectors for the cables ( 4 ) and ( 5 ). Otherwise the high voltage can create sparks.

- The cables ( 4 ) and ( 5 ) should be unscrewed from the transformer and placed

away from it.

- A new analysis should be started. When "GEN" appears and the fuse is not damaged,  and the main current didn’t exceed 10 A ( RMS ), then the transformer is functioning properly.

- Otherwise, replace the transformer. The cables ( 4 ) and ( 5 ) are then reconnected.

The cable ( 7 ) is unscrewed and removed from the capacitor ( 8 )

WARNING !  HIGH  VOLTAGE !

After the generator starts ( GEN ) the cable ( 7 ) carries a voltage of about 7000 Volts ! The cable ( 7 ) therefore should be laid so that its end is at least 10 cm away from the nearest object!

- If, after pressing < START > and GEN, the fuse is blown, then replace the rectifier ( 9 )

Further causes could be found in the oscillator circuit:

Because of the high voltage ( 7000 volts ) and the high frequency ( 19.5 MHz ), it is difficult to make measurements without any special measuring instruments.

Thus, the analyser needs to be turned off, and the power plug disconnected.

WARNING !   HIGH   VOLTAGE !

- Check that all the screws and soldering connections

have a good contact.

If the resistor 100 KΩ / 50 W ( 11 ) is a short circuit,

then the fuse will blow when starting the combustion

- Check that all power conductors and copper bands are

placed far enough away, from any nearby objects! The

wire ( 1 ) and the copper band ( 2 ), for example,

should be placed far away from the plate ( 3 ).

- The power cable ( 4 ) should have a maximum distance

from any surrounding object.

- The copper band ( 5 ) should be placed as far as

possible away, from the capacitor ( 6 ) !

This is why the copper band ( 5 ) should always

be bent around the insulator ( 7 ) !

If the tube needs to be replaced ( 8 )

- Loosen the screw ( 9 ), pull out the heat sink

( 10 ) upwards ! Pull out the tube ( 8 ) upwards!

## 1.16   The fuse blows at the end of the analysis:

Causes :

- The triac is malfunctioning.

The so-called Latch-up effect causes the undesired conduction during one halfway

of the triac, once the trigger pulses to the triac have subsided. The transformer is then

acting as a short circuit because it is fed with a DC voltage!

Replace the TRIAC!

( To locate the triac, see 4.1-3 )

## 1.17   Large collection of dust when opening the furnace:

After pressing the key to open

the furnace, the valves ( V1 )

and ( V2 ) a re-activated,

this means that on the coils 24 volts

DC coming from the electronic unit

are present.

Valve ( V2 ) closes the oxygen

supply to the furnace.

Valve ( V1 ) opens,

releasing the pressure from the furnace

through the dust trap ( 3 ),

through the moisture trap ( 4 )

and through the filter cartridge ( 1 )

into the open air

( in the direction of the arrow ).

Following this, the pneumatic valve shuts

and the furnace opens

( lift downwards ).

When dust comes out of the furnace after

this process, this will cause that the pressure

in the furnace is not sufficiently eliminated.

The causes:

- Too much dust in the dust trap ( 3 ), ( a few hundred analyses without cleaning )

- Saturated moisture trap ( 4 )

- Clogged paper filter in the moisture trap ( 4 ), see Operation Manual 6.1

- Clogged filter cartridge ( 1 and 2 ) (this would be unlikely though)

- One of the valves ( V1 ) or ( V2 ) receives no power ( 24 V DC )

CAUTION !

- The valve ( V1 ) will only be supplied with voltage ( 24 V DC) until the lift has reached the bottom position, then it will be switched off.

- The 24V on the coil of this valve can be measured with a digital voltmeter, or with a light bulb of 24 V / 20 mA.

- One of the valves ( V1 ) or ( V2 ) is sticking,  preferable valve ( V1 )

- Clean the valves, see 3.13

## 1.18   No or low oxygen pressure:

Induction furnace operation

SYMPTOM: After pressing the START button, the following message appears:

"no oxygen pressure "

POSSIBLE FAULTS: on the front panel of the analyser:

- Is the power switch ( 1 ) set on pos. 2 ?

- The pressure on the oxygen gauge ( 2 )

may be too low.

1.5 bar is normal ( 22 psi ).

The pressure remains constant when

the furnace is closed, and it sinks

during the upward movement

of the furnace floor.

- Check the oxygen pressure

in the supply line or tank.

4 to 6 bar ( 60 to 90 psi ) showing on

the pressure gauge is normal.

- Check if the pressure gauge for

compressed air ( 5 ) shows 4 to 6 bar

( 60 to 90 psi ).

- Check if the cock ( 3 ) is properly shut,

or else the O-ring ( 4 ) will not seal.

Check if the O-ring ( 4 ) is in a good condition.

- Check if the combustion tube is broken

or cracked. See Operation Manual 6.6

- Check for gas leakage, see 3.7

- Check if the chemicals are saturated

and have become hard.

- The paper filter could also be the cause.

For changing the paper filter,

see Operation Manual 6.3

On the rear-side of the analyser:

- The contents of the gas cleaning furnace could be worn out and clogged.

See Operation Manual 7.3

Inside the analyser:

The oxygen solenoid valve ( V5 ) may be defect.

- Open the left-side panel of the analyser.

Unscrew the upper tube ( 1 ). The oxygen should

rush out when the power switch is set to pos. 2

If not:

- There may be no oxygen supplied.

- 220 V AC is not measured on the terminal ( 2 )

of the solenoid coil.

- The solenoid valve is clogged.

Clean the solenoid valve, see 3.13.

If the fault is still present, then:

- Switch off the power, ( set the power switch to pos. 0 )

Pull out the socket ( 2 ) from the coil and  measure the

coil resistance. 3.3 KΩ is normal.

The outflow valve does not close :

Open the left-side panel of the analyser. When the furnace is closed ( the lift is up ),

no oxygen should be flowing out of the lower ( free ) opening of the valve ( 1 ) ( in the

direction of the arrow ). This can be checked by putting your finger under the lower

opening of the valve. If the valve does not close correctly, you will feel a strong

oxygen pressure. The power supply to the valve ( 1 ) should be 0 Volt. If this is the

case then the outflow valve ( 1 ) should be cleaned see 3.13, or replaced.

If the voltage on the outflow valve is constantly at 24 V, then:

- Erase the memory, see 3. 9.

- Replace the electronic driver, see 1.6.

OTHER FAULTS:

The screw for the upper furnace cushioning is too tight, see 2. 6

The pressure switch does not close, see 2. 2

The furnace closes too slowly, adjust according to 2. 6

The furnace does not close: see 1. 7

## 1.19   Saved data are lost:

- Make sure that a 74LS 04 is present.

- In this location only an LS-Type has to be installed.

- Replace the accumulator ( accu in diagram ), when the accumulator will not charge

after the power had been switched on, and after installing the 74 LS 04 .

Once the analyser is switched off and the accu didn’t charge data will be lost !!!

The data that is meant here is the data that is saved in the accumulator-buffered RAM area. It consists mainly of the calibration factors, the blank values, and the linearity factors, which are entered through the keyboard. See 2. 9

The linearity factors which were programmed in our company when the device was manufactured, were saved in the EPROM, they don’t get lost.

After turning on the power again ( with an empty accumulator ), those default linearity factors that are saved in the EPROM’s are the current values.  Therefore, you should note down any linearity factors that you enter through the keyboard, in case the accumulator becomes empty.

## 1.20   Noise or drift of the infrared baseline:

The symptoms are noticeable on the display and especially on the monitor

attached to the PC, when either rhythmic or sporadic jumps of the zero line are

present.

- The cause for a swinging rhythmic is the chopper motor.

Replace the chopper motor, see 3. 5

Causes for noise are :

- The pre-amplifier,

replace, see 3. 1 and 2. 3

- The infrared detector,

replace, see 3. 1 and 2. 3

- A fault on the circuit board IR 21,

replace, see 3. 3

- The power supply fluctuates,

see 3. 3 and 3. 4

- The IR source is defective,

see 3. 2 and 1. 8.

## 1.21   The life time of the combustion tubes is too short:

Possible reasons for the induction furnace quartz tube:

- The sample weight or the weight of the accelerator is too high.

2.        Don’t use copper as accelerator.

3.       The combustion tube is too long, i.e. longer than 145 mm. In this case, after fixing the four wing nuts ( 3 ) and ( 9 ), the combustion tube is pressed between the upper furnace closing part ( 17 ) and the lower part ( 10 ). See Operation Manual 6.7

In this case, when the combustion tube heats up during combustion, it may break.

Check the length of the combustion tubes to be not longer than 145 mm.

4.        The upper furnace closing part ( 17 ) does not provide enough space for the    combustion tube. To check this, install a combustion tube as usual but without the O-rings ( 11 ) and ( 12 ). After fixing the wing nuts ( 3 ) and ( 9 ), move the tube up and down. If the tube does not move at least 0,5 to 1 mm, replace the part ( 17 ) or make room for the combustion tube by machining the upper furnace closure on a lathe.

## 1.22   The word “ELTRA” appears on the display:

This only happens when the power supply to the electronic is turned on.

The electronic unit can function down to a minimum main voltage of only 160 V.

If, during the operation the word "ELTRA" appears on the display, then the

voltage has sunken below 160 V.

- Monitor the main voltage during an analysis.

Should there be a short interruption detected, then the cause could lie in a poor contact.

- If the voltage drops when the generator is running, then the generator power can be reduced:

The cable that is connected on the high voltage transformer to the 220 V terminal,

should be then connected to the 270 V terminal. This way, the generator draws

less energy.

- Use a voltage stabiliser, see 3.15

## 1.23   The oxygen pressure falls when the induction furnace opens:

Regardless of whether the furnace is

opened or closed,

the oxygen pressure shown on the pressure

gauge ( 2 ) must remain 1.5 bar ( 22.5 psi )

constant.

If, however, the oxygen pressure falls

when the furnace is open, then:

Either the valve ( V2 ) is defective or its control is incorrect.

When the furnace is open, the valve ( V2 ) should have 23 V DC on its magnetic coil.

When the furnace is closed, the voltage should be at zero volts.

- If the voltage on the valve ( V2 ) remains at 23 V DC, no matter whether the furnace is opened or closed, then the electronic unit is defect, see 1. 6.

- If the valve voltage always stays at zero, then either the wiring to the electronic unit is broken, see 4. 2; or the electronic unit is defect, see 1. 6.

- If the voltage on the valve ( V2 ) is correct, then the valve itself is defect, see 3.13.

## 1.24   Combustion without starting the analysis:

Induction furnace operation

In this case, the combustion of the sample begins immediately after the sample enters the furnace, and before having pressed the "START" button.

Should the combustion start just

after closing the furnace, and by pressing

then there is a serious fault !

The power switch should be set immediately

to pos. 1, in order to avoid any overheating

of the generator.

- The electronic unit ( rack ) is out of order, when the combustion starts without pressing "START" and without "GEN" on the display, there is only 0,5 V DC on TP9     ( instead of 12 V DC ). Replace the electronic drivers, see 1. 6.

When the pulses at TP13 are present without pressing "START"

and with 15 V DC at TP9, then the circuit board HF 42 is out of order.

If no scope is available, the pulses can be detected by using a voltmeter.

Normal values are about 0 V AC without pulses, and about 1 V AC with pulses.

This indicated value can be different, depending on the type of the voltmeter used.

Possible reason :

- The triac is faulty ( short circuit ), and has to be replaced.

( To locate the triac, see 4.1- 3 )

## 1.25   The crucible cracks or melts:

One reason could be that:

The quality of the crucible is bad, use a better quality.

Another reason could be that:

The furnace temperature is too high, due to:

- Excessive main voltage.

- Excessive sample and / or accelerator weight.

- The use of a “sample or accelerator” that requires lower generation power.

In this case reduce the generator power, see 2.12

## 1.26   The analysis doesn’t start :

There are mainly two reasons for the above fault:

- Insufficient oxygen pressure ( below 1 bar ), see 1.18

- The infrared zero baseline deviates too much from zero, see 2.3

## 1.27   The oxygen pressure goes up very slowly:

When using induction furnace operation:

After the button has been pressed to close the furnace ( or to start the analysis by

closing the furnace first ), the oxygen gauge ( 6 ) should reach the normal level of

1,5 bar within a couple of seconds after the furnace has closed.

If this takes longer:

Then the problem is outside the analyser.

- Either the plastic oxygen tube is squashed from the analyser,

or the oxygen regulator is not properly open.

## 1.28   Pump capacity vs maximum:

For resistance furnace operation

Faults can easily be located with the help of this flow diagram.

It is possible that the above error message, “too high voltage on the pump”, is triggered not only by one, but by a number of faults.

A typical example could be the combined partial blockage of the glass tube ( 22 ), and of the

combustion tube.

Always ensure, that the chemicals are not saturated and that the combustion tube is not blocked.

Usually, these faults caused by contamination occur due to frequent analysing with samples that produce a lot of dust, and after a prolonged use of the combustion tube.

17	Pressure gauge 1,5 bar			35	Furnace inlet - oxygen

20	Carrier gas flow gauge			36	Furnace outlet measuring gas

21	Measuring gas flow gauge

22	First measuring gas purifier

23	CO2 / H2O - trap

24	Catalysis furnace

25	SO3 - trap

26	H2O - trap

27	Dust filter

1	Carrier gas inlet		14	Solenoid valve

2	Pressure regulator		15	Lance valve

3	Pressure regulator		16	Pressure switch

4	Pressure regulator		18	Restrictor

5	Oxygen valve		19	Restrictor

6	Bypass valve			28	IR-cell

7	Flow regulating valve	29	Gas pump

8	Purge valve			30	Attenuate before the pump

9	Exhaust valve		31	Attenuate after the pump

10	Outflow valve		32	Furnace

11	Solenoid valve		33	Induction furnace

12	Solenoid valve		34	Gas cartridge

13	Solenoid valve		37	Flow sensor

Detailed explanation of the faults indicated in the flow diagram:

Blockage of the lower connection of the glass tube ( 22 ) located between the

power switch and gas flow gauge:

The bore hole of the lower connection of the glass tube ( 22 ) can get blocked when there isn’t enough quartz wool at the bottom end of the glass tube ( 22 ). This can cause that fine particles of the magnesium perchlorate enter the bore hole therefore blocking it. It is therefore absolutely important to fill the glass tubes exactly according to the Instructions. See Operation Manual 6.3

The bore holes need to be cleaned from both the inside and the outside of the analyser,

due to the fact that the passage is not straight, but bent at 90°.

CAUTION!

Switch off the main power before unscrewing the tube from the connection.

If the connection is blocked, then the tube may be blocked as well.

Therefore the tube needs to be replaced, or else the contamination could reach

the IR-cell as well.

The IR-cell is blocked:

- This fault is caused mainly by a kink in one of the tubes.

A blockage due through contamination is practically impossible. It is more likely

to find that the lower connection of the glass tube ( 22 ) gets blocked.

Leakage:

- E.g. from the attenuator ( 30 ), or from the gas flow meter ( 21 ):

The message „Pump capacity vs. maximum“ is not only triggered through blockages,

but also by leakages.

The leakage is located between the flow sensor ( 37 ) and the gas flow pump ( 29 ).

Occasionally , the cover of the attenuator ( 30 ) is not properly sealed, or the gas flow

meter ( 21 ) has a crack on the joints for the tube connections. Disconnect the gas flow

gauge and check for cracks.

Blockage in the lower connecton of the thick glass tube ( 31 ):

As already mentioned the lower connection of the glass tube ( 30 ), can get blocked

if there is not enough quartz wool used.

- Remove the glass tube ( 31 ).

- Remove the tube from the lower connection.

- Clean the bore hole.

- Another possibility could be that the cotton wool inside the glass ( 30 ) of the attenuator prior to the pump is very dirty.

- CAUTION!   Please follow the correct sequence of the glasses:

- without cotton wool         with cotton wool

- The end of the tube

- in the glass must be

- free. It shouldn’t be

- blocked by cotton

- wool. The same

applies for the

fitting at the outlet

of the glass.

- The pump is joined to the short connection inside the glass.

Furnace:

- Clean according to Operation Manual 6.4

## 1.29   Resistance furnace takes too long to warm up:

Normal warm-up time: from room temperature to 1300°C, in 15-20 min.

Faulty heating elements:

Investigate the problem in the following way:

- Set the demand value to maximum, measure the voltage across the heating elements.

- Set the measuring range of the voltmeter to 1000 V AC.

If the voltage is near 220 V AC, then the heating elements are faulty.

- Replace the heating elements, see 3.18

- If the measured voltage is much lower than 220 V AC, then the set value for the

current limitation is too low, see 2. 6

## 1.30   Negative drift after the start:

This problem is caused when some material from previous samples, fall off the combustion boat when shifted into the furnace.

Then, during the extraction from the furnace, the hot combustion boat may touch and ignite this material, causing an unwanted small peak of the base line.

If the following analysis is started while the base line is still above zero, then the electronics will tare that momentary level.

After the start, the base line will drift negatively, because the peak of the burned particle is descending towards the original zero.

A problem of the IR-cell itself can be ruled out, when running a few analyses by just entering the weight into the memory and pressing “START”, without inserting or extracting the combustion boat. If the base line doesn’t show a negative drift, then the IR-cell is O.K.

## 1.31   The resistance furnace is completely cold or too hot:

## 1.32   Smoke in the resistance furnace:

- It is a sign of bad combustion. See 1. 5

- The furnace temperature is too low. See 1.14 and 1.17

- No oxygen flow. See Operation Manual 2. 4

CAUTION !

In this situation, never turn on the oxygen flow, or else there can be an

EXPLOSION !

Wait for at least 15 minutes, until the smoke has disappeared from the furnace.

- The pressure switch is either unadjusted or faulty, see 2. 2

- The oxygen valve is faulty, check for pressure response.

Unscrew the lower tube ( 1 ).

The oxygen must exit the valve when the main switch is set to pos. 2

If not:

- There is no voltage ( 220 V ) on the connections ( 2 ) of the magnet coil ( 4 )

- The valve is blocked. Clean the valve by:

- closing the oxygen supply,

- setting the mains switch to pos. 0

- clean the valve according to 3.13

If the fault is not removed, then:

- Switch the power off ( pos. 0  )

- Disconnect plug ( 2 ) from the magnet coil ( 4 ) and measure the resistance.

3.3 K is normal.

## 1.33   Resistance furnace doesn’t reach full temperature:

- The heater elements are used up. See 3.18

- The furnace is faulty. See 1.31

## 1.34   Unstable resistance furnace temperature:

- The heater elements are used up, and the required temperature is delivered with great difficulty. There are no reserves of power, to keep the temperature stable. See 1.33

To confirm this fault, set the temperature to a much lower value. The required

temperature should then stay constant.

- The temperature control board TH 43 is faulty.

Replace the TH 43 board and adjust according to 2.15

- The thermocouple is faulty. See 2.15

## 1.35   Dust or soot from the resistance furnace:

- A sample that has been inserted in the furnace too quickly.

See Operation Manual 3. 4.

The sample weight is too high:

- If a sample burns explosively, then there could be a shortage of oxygen.

This can cause some soot or oil vapours that  passes through the hot zone of the

furnace without burning.

It is necessary to reduce the weight of the sample, e.g. to 100mg or even to 50mg.

If the weight reduction doesn’t help, then operate differently by shifting the boat

slowly towards the beginning of the hot zone of the furnace. Leave the boat there

and allow the sample to start burning. Towards the end of the combustion, move the

boat to the middle of the hot zone, in order to obtain a complete combustion. See

Operation Manual 3. 4

The sample is metallic:

- perform analysis according to Operation Manual 3. 4

## 1.36   Vibrations in the induction furnace mode:

The reason is dust accumulation on the blades of the oscillator’s blower. The oscillation results due to distribution of dust on the blades causing an unequal rotation.

Caution:

Disconnect the power before

Removing the oscillators lid.

Blow the dust away with compressed air. Blow from the inside to the outside, otherwise the dust will be blown into the oscillator. Take care of the rotary speed of the blower to remain within the operation limits. When blowing compressed air, the rotary speed of the blower can exceed the limit of operation causing damage to the bearings and the shaft.

## The analysis cycle stops only 25 seconds after start:

Information.

When 50 seconds as maximum analysis time is entered, this means that 50 seconds after start the analysis resp. the combustion will definitely stop in any case. But if there is no signal ( peak ) coming from the infrared cell, the analysis cycle will take at least 25 seconds i.e. half the maximum analysis time.

The problem.

It seems that there is no signal coming from the infrared detector. If the peak comes the minimum analysis time of 25 seconds will be automatically prolonged until the signal comes back to zero or the maximum analysis time of 50 seconds has been reached.

There are two possibilities for not having a signal:

- There is no combustion. See 1. 4 and 1. 5.

- The infrared cell is damaged. Do the IR-test. See 1. 1, 2. 3 and 2. 4.

Make sure that you take a sample with relatively high carbon to provide a clear visible high peak to about 3 to 8 volts.

## The chopper motor doesn`t rotate:

Mostly, the reason is that the grease of the bearings becomes dry.

You can feel a resistance when you try to turn the shaft by hand.

In this case the problem can be solved by lubricating using silicon oil.

Dip a piece of wire in to a flask of silicon oil and apply as shown below.

In order to spread the oil into the space between shaft and bearings, turn the shaft and move it in axial direction at the same time. The use of mineral oil is not suggested, because mineral oils become dry and hard after long time of operation. If no silicon oil available, use silicon spray. Silicon grease is also good for this purpose and it is also available in laboratories, however it could be very difficult for the grease to penetrate the space between shaft and bearings.

## High carbon results are erratic:

If only high carbon results are erratic, while low carbon and sulfur results are repeatable, most likely the metal dust filter is blocked.

Even if the filter may be brushed and cleaned in ultrasonic cleaner, some day will not be able to be cleaned any more. The filter doesn’t last for ever.

Check by removing the filter with the dust filter housing.

Don’t remove the filter from its housing.

Turn the filter housing upside down and try to blow air through the small hole (about 16mm diameter) of the filter housing.

If you feel a remarkable restriction, replace it by a new one. With a new filter the air should pass practically free without any restriction.

# 2 ADJUSTMENTS, TEST AND FUNCTION EXPLANATION

## 2.1   Gas flow controller adjustment and jumper settings:

- Make sure that the pressure regulators are at least pre-adjusted. See 2.7. This is mainly important for the pressure regulator ( PR3 ) in front of the gas flow regulating valve.

- Set the flow rate with P1 of the HF 42, to about 180 l/h. Observe the flow rate on the lower flow meter of the analyser, wait until the value has settled.

- The output of the flow sensor is found on TP10 of the HF 42 board; TP15 is GND.

With a flow of 180 l/h, the voltage is about 2.5 – 3 V.

Please note: without flow, the output voltage of the sensor is not 0 V, but about 0.6 V.

- Adjust with P3 until the voltage on TP7 is 5 V.

- On TP8, the activity of the pulse width modulator can be checked.

- The output voltage to the gas flow regulation valve can be measured on TP3.

With a constant flow of about 180 l/h, it should be between 5 and 15 volts.

The panel meter ( 4 ) must indicate between 30 and 70 µA.

- Check and if necessary readjust the pressure regulator ( PR3 ) in front of the flow regulating valve. See 2.7

Jumper settings:

Detailed explanation of the gas flow regulation:

The flow controller mainly consists of the following components:

- The flow sensor ( 1 )

- The flow regulating valve ( V6 )

- The flow control board HF 42

The flow sensor detects the gas flow, and converts its amount into an electrical voltage.

This voltage is then sent, as a feedback signal, to the HF 42 board.

The HF 42 board compares this feedback value, to the flow set point value.

The signal resulting from the comparator, runs a pulse width modulator.

Its output can be seen on TP8.

The pulse width modulated signal runs a power transistor, which in turn, supplies the output signal to the flow regulating valve.

The duty cycle of the output signal determines its DC voltage, which, with a flow of about 180 l/h and with clean chemicals, measures around 9 V ( on TP3 ).

The HF 42 board drives the flow regulating valve ( V6 ) in such way, to provide a constant and accurate gas flow. The µA meter ( 4 ) shows the rate, at which the flow valve ( V6 ) is driven „open“.

Remark:

Due to a few failures of the flow sensors, almost exclusively on the CS-2000 at customers who regularly treat their samples with acid, we concluded that the reason of the failures should be of chemical nature i.e. acids and Schutze`s Reagent damage the sensor.

The solution is to prevent the analysis gases to get in contact with the sensors. Therefore the sensor ports are connected with long tubes to the gas flow system. There is no gas flow through the sensor tubes but only the pressure is sensed by the sensor.

In case of failure of a previous version sensor, it is advisable to replace it by the new version. In this case the part numbers of the complete assemblies should be used for ordering. In case the analyser has already the new version sensor, the order should include the spare part number. See attached drawings.

The new flow sensor can not be used as a replacement for the older flow sensor consisting of a rotameter and a optical element.

## 2.2   Pressure switch:

Induction furnace operation

- Close the furnace.

The pressure switch should be adjusted to about 1 bar ( 15 psi ); this means that under

1 bar ( 15 psi ), the contact is open ( no contact ), and over 1 bar ( 15 psi ), the contact

is closed ( turned on ).

- In installed condition and with the analyser turned on ( Power switch set to pos. 2 ), there should be a voltage of 24 V DC when the pressure is under 1 bar ( 15 psi ).

If the pressure is over 1 bar ( 15 psi ), then the voltage should be zero.

ATTENTION !	Measure this with a high-impedance digital voltmeter!

Do not cause a short circuit while measuring, or else the

electronic drivers will blow

- The oxygen pressure can be adjusted to 1 bar ( 15 psi ) by turning the inlet pressure regulator on the rear panel of the analyser. See 2. 7. This can be read off the upper

pressure gauge of the analyser.

- The switch point can be set by turning the screw ( 1 )

- If the switch is turned off ( 24 V ) when the pressure is at 1 bar ( 15 psi ), then the screw ( 1 ) should be turned counter clockwise until the switch turns on, measuring the power until it reaches zero volt.

- If zero volts are measured when the pressure is at 1 bar ( 15 psi ), the screw ( 1 ) should be turned clockwise until the voltage goes to 24 V.

Resistance furnace operation:

Set to resistance furnace operation ( power switch position 2 ).

The pressure switch for resistance furnace operation is adjusted in a similar way like the pressure switch for the induction furnace operation.

CAUTION:  Select 750 VAC range. The voltage of this pressure switch is 230 V AC.

## 2.3   Infrared zero- baseline, control and adjustment:

The infrared cell zero-baseline is normally controlled by the electronic and automatically set to zero.

The auto – zero control should be disabled In order to control and adjust the infrared cell zero-baseline manually, this way the electronics will have no influence. This can be done by pressing the   < Menu >  and  < Cal >  keys.

The voltage exiting the infrared channel should be shown as in the following example:

- 0.13

- 0.21

- 0.18

- 0,17

The displayed numbers refer to the IR-cell output voltage:

The upper number refers to the voltage of channel	1

The next number refers to the voltage of channel	2

The next number refers to the voltage of channel	3

The lower number refers to the voltage of channel	4

The channel numbers can be shown on the display by pressing:  < Menu >  < 5 >

Return by pressing  < CLR >.

If at least one voltage exceeds +3.0 volts, then the zero-signal needs to be adjusted.

- To adjust the zero signal, push the left cover plate ( 2 ) of the infrared cell about 10 cm towards the back.

With a small screwdriver, adjust the trimmers ( 3 / 4 / 5 / 6 ) in the infrared rack, so that the numbers shown on the display are between + 0.20 and - 0.20.

The adjustment should preferably be carried out while the carrier gas is flowing.

If the zero level ( between + 0.20 and - 0.20 volts ) cannot be reached, and if more

than two adjustments were recently carried out, then the problem could be:

- Contamination of the IR-paths, see 3. 1

- bending of the IR-source, after rough shipment of the device, see 3. 2

- insufficient supply voltage for the Infrared, see 2. 4

The stability of the displayed numbers should be better than ±0.01.

Normally, the second digit after the decimal comma does not change, or it may

fluctuate up and down by only 1, since this could be actually on the rounding point.

If the fluctuations are larger, then there is a major degree of noise, see 1.20

The drift should also be very small. Within 10 minutes the second digit after the

decimal point should not change more than 1.

The drift should first  be checked when the analyser has been switched on

( power switch on setting 1 ) for at least one hour, better would be two hours, and

the oxygen flowed for about 10 to 15 minutes ( Power switch on setting 2 ).

In case the drift is too large :

- Check the temperature regulation, see 2. 5

- Unstable power supply, see 2. 4

- Defective IR-source, see 3. 2

- By pressing the < CLR > key, you can end the test.

- Notice :

If the infrared cell's output voltage is over 10 volts,

++++ is shown on the display, ++++ then there is a serious defect.

Adjusting the infrared cell:

The connection sockets of the measurement channels ( 9 / 10 / 11 / 12 ) are predetermined on the circuit board IR 41.

The position of the IR-paths on the path plate ( 25 ) is irrelevant. Each IR-path can be changed, according to the required measuring ranges.

The connections to the IR 41 board, however, is very important:

9		Connection for section ( 1 )

10		Connection for section ( 2 )

11		Connection for section ( 3 )

12		Connection for section ( 4 )

( C or S, depending on the layout of the Analyser, see 3.11 )

13 / 14 / 15 / 16		Pre-amplifiers

17 / 18 / 19 / 20		Adjusters

21 / 22 / 23 / 24		Zero fine adjustment

- Each channel is adjusted separately.

- Set the trimmers ( 19 / 20 / 21 / 22 ) in the middle of their range.

- Connect a digital voltmeter on the outlet 6 and the ground 7

( measuring range 200 V DC )

- Adjust 0 volts DC by adjusting the preamplifiers (13 / 14 / 15 / 16 )

- Turn the trimmer clockwise for lower voltage, counter clockwise for higher voltage.

After the zero adjustment, the voltage on test point 1 should be about 1 to 1.5 volts

AC ( measuring range 2 V AC )

- Install the infrared cell in the Analyser.

- Restore the oxygen flow.

- After 30 to 60 minutes, adjust the zero by turning the trimmers ( 19 / 20 / 21 /22 ).

- To adjust zero, the  < Menu >  and  < Cal >  keys are pressed,

- Press the  < Cal >  key again to return to the analysis.

- On the temperature regulation card, check if the LED is blinking, see 2. 5

ATTENTION !

When the test signal fails on test point 1, this can be due to the 8-pin connector

( 9 / 10 / 11 / 12 ) that connects the pre-amplifier to the IR 41 board not being correctly connected.

- remove, turn 180°, and replace, the 8-pinn connector and measure again.

1	black		marked tube connection

2	yellow	marked tube connection

3	red		marked tube connection

- When reinstalling the infrared cell into the analyser, you must pay attention to the colour markings of the tubes and their connections.

.

## 2.4   IR-source voltage setting:

IR-source voltage, is to

be measured between TP1

and TP2 on IR 45 board.			Transformer connections:

Blue cables

1,75 V DC   for  1 channel device		1  Soldering Location for 1 channel device

3,5   V DC   for  2 channel device		2  Soldering Location for 2 channel device

5,25 V DC   for  3 channel device		3  Soldering Location for 3 channel device

7      V DC   for  4 channel device		4  Soldering Location for 4 channel device

( 1,75 Volts DC for each IR source )	8  Always connected ( common )

- Set above voltages through Jumpers

- The voltage of the capacitor C must be about 5 to 6 V higher than the voltage of the IR-source measured on the a.m. test point.

- If the voltage of C is too low, either the transformer is not correctly connected, or the rectifier R is out of order.

## 2.5   Temperature regulation of the IR-cell:

The adjustment should preferably be done only when the analyser, and the

infrared cell, has been in operation for at least one hour.

- If the LED is always off, adjust  P1 on the TH 62 board clockwise for higher

temperature set point.

- If the LED is always on, adjust P1 counter clockwise for lower temperature

set point.

In case the LED is already blinking, adjust the flash rate of the LED to 50%

duty circle.

ATTENTION!

- The effect of the readjustment will appear in a few minutes, as soon as the temperature has changed to the new setting.

- The voltage on TP1 ( TP2 = GND ) does not correspond to the actual temperature of

the IR-cell, but it corresponds to the power needed to keep the IR-cell at a constant

temperature.

- When the environmental temperature goes down, the heaters of the IR-cell need

higher power to keep the temperature of the IR-cell constant. In this case the

voltage will go higher

The opposite happens when the temperature in the laboratory goes higher. The IR-cell heaters need less power to keep the temperature constant, so that the voltage will become lower.

- The adjustment of the IR-cell should be done, if possible, when the laboratory has the average temperature of a year. The set point of the IR-temperature should then be adjusted with P1 so that the voltage in stable condition measures 6 volts.

- For readjusting, turn P1 when the voltage too is to low clockwise, when the voltage is to high turn counter clockwise.

- Caution: After you have turned P1, several times the voltage will be out of range. This is normal. When turning P1 you enter a new set point of the IR-cell temperature which is at that moment different to the actual temperature of the cell.

- The temperature control circuit will give full power to the heaters ( in case you have set a higher temperature ) or the circuit will completely shut down the power to the heaters, if you have adjusted a lower temperature set point than the actual temperature.

- After about 10 minutes the temperature will reach the new adjusted level, and the voltage will come back to remain stable at the new level.

- This is the voltage level which corresponds to the power to keep the

IR-temperature stable at the new level.

- If the LED still doesn’t start blinking , adjust P1 on the TH 62, until there is about

5.35 V DC present between TP2 an TP3.

- If the fault still remains, despite adjusting the TH 62, then check the board NR 32,

according to the following instructions.

- CAUTION!  ELECTRICIAL  HAZARD !

- You should be able to measure 200 V DC between TP1 and TP2 ( GND ) of the NR 32 board; otherwise the board is faulty.

- This requires that 220 V AC comes through the connectors 6 and 7

- If there is no 220 V AC, then there could be a bad contact  in the plug.

- The voltage for the infrared cell heating elements, comes from the connections 4 and 5 of the 7 pin plug, measure the voltage on the screws, without disconnecting the plug.

- As long as the LED is on, there is 200 V DC between 4 and 5.

- As long as the LED is off, there is no voltage between 4 and 5.

If there is no voltage between 4 and 5, then the NR 32 is faulty. Replace it, according to 3. 6.

- The above sentences, extend to the board TH 61. From there they feed the heating

elements.

General description of the infrared cell temperature regulation:

For more details see 4.2

Simplified diagram with TH-62

This system consists mainly of the following components:

- The temperature sensor

- The temperature control board TH 62

- The power stage board NR 32

- The heating elements

The temperature sensor feels the temperature of the infrared system, and sends the signal to the temperature control board TH 62.

The TH 62 compares this signal with a set point value. The resulting value drives a pulse width modulator. The LED of this board shows the function of the modulator.

The modulated signal further drives and power stage of the NR 32 board; the power signal returns to the TH 62 board, from there it finally goes to the heater elements.

## 2.6   Pneumatics:

1  Combustion coil			  6  T-piece			11  Pressure gauge

2  Muffler				  7  Lower cushioning	12  T-piece

3  Upward speed regulator		  8  Pneumatic cylinder	13  Pneumatic valve

4  Muffler				  9  Magnetic switch		14  Coil cooling valve

5  Downward speed regulator	10  Upper cushioning	15  Pressure reduction valve

PR4  Pressure regulator

Adjusting the pneumatics for correct furnace movement:

- Check if the air pressure on gauge ( 11 ) is between 4-6 bar.

- Press the button  ↓  ↑  to open the furnace.

- Open the right-side panel of the analyser, to find the pressure regulator PR4 and the pneumatic valve ( 13 )

- Turn the screw ( 5 ) counter clockwise, for minimum air restriction. Leave it in that position.

- Rotate the knob of the pressure regulator PR4 fully counter clockwise.

- Turn the screw ( 3 ) clockwise, for maximum air restriction. Leave it in that position.

- Shift the magnetic switch ( 9 ) fully upwards.

- Press the button  ↓  ↑  to close the furnace ( furnace floor won’t  yet move ).

- Adjust regulator PR4, until the furnace floor can rise completely, without shaking,

- For safety precaution, a person still can stop moving the furnace floor with only one hand.

- When the furnace is closed, shift the magnetic switch downwards until it switches on;

Now shift down another 5 mm.

- Open the furnace and then close it, to check if the adjustments are still correct.

- Adjust the furnace upwards speed with screw ( 3 ); it must be within a reasonable limit.

- Adjust the furnace downwards speed with screw ( 5 ); it must be within a reasonable limit.

- Adjust lower cushioning ( 7 ), so that the furnace opens without banging at the bottom.

Check the following :

- Open the furnace ( piston down ).

- Hold the piston down by pressing on the horizontal

bar of the furnace. ( with the fingers )

- Press the key to close the furnace.

- Allow the piston to go up very slowly by controlling

the force of your fingers.

If not possible do the pneumatics adjustment.

see pages before.

- About 5 mm before the piston reaches it’s highest

position the force will increase.

- With your fingers press down firmly to make

sure that the piston can not be pushed down.

- If the piston starts vibrating by moving up and

down with full force, then adjust the switch higher.

- If the piston goes down, then adjust the switch lower.

## 2.7   Pressure regulators:

PR1 is the inlet oxygen regulator, and is situated inside the analyser, on the rear panel.

The outlet pressure needs to be set to 1.5 bar ( 22.5 psi ).

- Pull and rotate the knob of the regulator, until the pressure gauge ( 6 ) shows the correct value.

- Ensure that the flow shown on the flow meter ( 5 ) shows between 180 and 200 l/h; if not, adjust according to 2. 1.

PR2 is the pressure regulator for the purging system of the IR-cell and it is situated

inside the analyser, just behind the left-side of the panel. It is positioned horizontally.

The oxygen pressure at the outlet of the PR2 should be 0.35 bar ( 5 psi ).

- The pressure can be checked by applying a T-piece on the outlet of the pressure regulator, and connecting  a pressure gauge ( 1 ), as indicated above.

- Pull and rotate the knob of the pressure regulator ( PR2 ), until the correct pressure is shown on the gauge.

Using induction furnace operation:

PR3 regulates the pressure of the carrier gas that circulates through the IR-cell and

then finally out of the analyser. This pressure regulator is situated inside the analyser,

just before the flow regulation valve ( 2 ). The gas pressure at the outlet should be

0.7 bar ( 10 psi ).

- The pressure can be checked, by applying a T-piece on the outlet of the pressure regulator, and connecting a pressure gauge ( 1 ), as indicated above.

- Pull and rotate the knob of the pressure regulator ( PR3 ), until the correct pressure is shown on the gauge.

- Ensure that the flow shown on flow meter ( 5 ), see drawing on previous page, shows between 180 and 200 l/h; if not, adjust according to 2. 1.

For more accurate adjustment do this:

Increase the pressure slowly by turning clockwise and observe the upper pressure gauge  ( 6 ) on the front panel. It may start vibrating. It may not vibrate even if you adjust the maximum possible pressure with ( PR3 ). The maximum possible pressure is slightly lower than 1,5 bar ( 22 psi ), because the incoming oxygen pressure is adjusted to 1,5 bar ( 22 psi ). After reaching the maximum pressure or when any vibrations should start, slowly reduce the pressure by turning PR3 counter clockwise.

Read the pressure on the gauge ( 1 ) when the vibration on ( 6 ) stops.

( In case of no vibrations read the maximum pressure you obtained ). Then go on reducing the pressure with ( PR3 ) and observe the volt meter ( 4 ). When it starts moving towards its maximum scale ( 100 µA ) stop reducing the pressure, and read it on the gauge ( 1 ).

Now adjust the pressure with ( PR3 ) at a level on the gauge ( 1 ) wich is about in the middle between the lowest pressure you read and the pressure you noted after the vibrations stopped.

For example:

The vibration stopped at the level of 1 bar. The volt meter ( 4 ) started moving to its maximum at 0,3 bar.

Therefore:      1 bar + 0,3 bar    = 0.65 bar on the gauge ( 1 ) with PR3

2

## 2.8   Closing cone, adjustment:

- You should ensure that the shaft ( 1 )

is securely fastened to the base ( 2 )

If not, hold the base ( 2 ) securely

and tighten the shaft ( 1 )

( engineers wrench size 17 )

- In order to adjust the cone,

the screws ( 4 ) should be loosened

so that the cone ( 3 ) moves horizontally

on the base ( 2 )

- Close the furnace ( with the lift upwards )

- Tighten the screws ( 4 )

- The cone is now positioned correctly

- The screws ( 5 ), ( 6 ), ( 7 ), must

now be well tightened in this order

- The screws ( 8 ) should be properly fastened

( with a wrench )

- The screws ( 9 ) should only be tightened

by hand

## 2.9   Linearisation:

The linearisation of the Analyser is done in our company and does not need to be repeated by the customer.

Therefore it is fatal to experiment with the linearity.

The following description applies only in case the number of channels are increased

see 3. 12 , or if the range of a channel has changed, ( for example, by installing a longer or shorter IR-path ).

The infrared cell is basically non-linear; meaning that its output voltage, is not proportional to the CO2 or SO2 concentration. In order to obtain the correct measurements, this non-linearity must be corrected electronically through the software. Once the „linearity factor“ has been entered, see description below.

For each measurement channel, the exact characteristic curve must be individually determined and programmed into memory.

After programming in a ( new ) factor,  analyse a high and a low standard sample, to check the accuracy.

If, the measurement results after the test of the low standard sample are not correct, a new factor must be programmed and the test repeated.

Procedure:

1 )	Analyse the high standard ( 7 Volts to 8 Volts peak)

2 )	Note the result

3 )	< Cal >  < 1 >  < Channel Number >

4 )	< Standard Value >  < Enter >

5 )	< Result >  < Enter >

6 )	Analyse the low standard ( 2 Volts to 3 Volts peak )

7 )	If the result ( from 6 ) is too low, then the factor should be higher.

8 )	If the result ( from 6 ) is too high, then the factor should be lower.

9 )	< Menu >  < . >  < Channel Number >

10 )	< New Factor >  < Enter >

11 )	< Nr >  < New Factor >  < Enter >

12 )	Repeat from  # 1

Approximate standard samples for linearisation and calibration of the

standard ranges for metal analysis with the Induction furnace:

Channel 1 :  Carbon ( low )  linearisation factor    Approximate value 23000:

0.197%C				400 mg

0.151%C				500 mg

0.100%C				500 mg

0.067%C				500 mg

0.023%C				500 mg

The standard for successive calibration is 0.151%C

Channel 2 :  Carbon        Linearisation factor           Approximate value 23500

3.920%C				    500 mg

3.200%C				    500 mg

2.080%C				    500 mg

1.050%C				    500 mg

0.452%C				    500 mg

0.243%C				    500 mg

The standard for successive calibration is, 2.080%C

Channel 3 :  Sulphur        Linearisation factor         Approximate value 50000

0.230%S				    500 mg

0.167%S				    500 mg

0.103%S				    500 mg

0.050%S				    500 mg

0.022%S				    500 mg

The standard for successive calibration, is 0.103%S

When the true factor is closely approached after using the above methods, an exact determination can be derived in the following way:

Rather than only one analysis, three separate analyses are conducted.

The arithmetic average is then taken.

1 )	Analyse the high standard three times.

2 )	Note the arithmetic average.

3 )	< Cal >  < 1 >  < Channel Number >

4 )	< Set  Value >  < Enter >

5 )	< Result >  < Enter >

6 )	Analyse the low standard three times.

7 )	If the arithmetic average ( from 6 ) is too low, then the factor should be higher.

8 )	If the arithmetic average ( from 6 ) is too high, then the factor should be lower.

9 )	< Menu >  < . >  < Channel Number >

10 )	< New Factor >  < Enter >

11 )	< Nr >  < New Factor >  < Enter >

12 )	Repeat from  # 1 )

Suggested materials for linearisation for resistance furnace operation:

Carbon range 10% C, linearisation factor approx. 23000

Calcium carbonate       12%C         300 mg

Calcium carbonate       12%C         100 mg

Carbon range 100%C, linearisation factor approx. 23000

Graphite                      100%C         500 mg

Calcium carbonate        12%C         500 mg

Sulphur range    2%S, linearisation factor approx.  65000

AR 1705                      1.49%C         450 mg

AR 1701                      0.50%S         450 mg

## 2.10   Programming the balance:

Basic settings for all types of balances:

2400 Baud

No parity

Send continuously

Please note:

We are not responsible for balances that have not been delivered through our company.

The below setting is only a general guideline. Some balance may have a different data or cable configuration and may therefore fail to communicate with the analyser. The balance

supplied by Eltra is a Sartorius, model CP 64. The balance is programmed ( always

with 3 Digits ) in our company, in order to establish a good communication with the analyser.

Should any problems arise, then please follow the next instructions:

At first switch the balance off by pressing                and follow the steps below.

1) Switch the balance on, by pressing the TARE button by holding it down. Then

press the on/off button.                 Release first the                 and then TARE.

2) Selecting the first programming  mode, 113 ( to stabilise the vibration ) by pressing

the first digit No. will appear on the display. Change the digit No. by pressing

Repeat step 2 until the No. 113 has been inserted then,

Store with a short press on the  TARE button.

3) Selecting the 2nd  programming mode 515 ( 2400 Baud )

Same usage as in step 2) and don’t forget to store with TARE ( short press )

4) Selecting the 3rd programming mode 614 ( continuous data sending )

Same usage as in step 2) store again with TARE ( short press )

5) After the programming has been done it is IMPORTANT to hold the TARE button again

until the display shows a row of 88888888 and then release the TARE button and the

balance is ready for operation.

The correct settings:

113   to stabilise the vibration

515   2400 baud

614   continuous data sending

CAUTION:

The balance model CP 64 has is a switch to enable / disable the programming mode.

The switch, when shifted to the left, disables the programming mode.

## 2.11

## 2.12   Induction generator control:

Current control:

The current control provides a safer and more accurate control of the furnace

intensity.

- Set the jumper J2 to B1-B2

- Connect a digital multi meter on TP1 ( TP15 is GND )

- Start the analysis, adjusting P2 until  6.3 volts are on TP1.

Adjust only, while the furnace is active.

The following diagrams show the relationship between the set value ( adjustment with P2 on the HF 42 board ) and the generator parameters.

## 2.13   Diagnostic printouts:

By pressing the:

< Menu >  < . >  < 0 >

keys, a diagnostic report will be printed out on the printer.

Attention !

The printer must be connected to the analyser.

The printer port must be activated, with :

< Menu >  < 3 >  < 1 >  < Enter >

This report should be faxed to ELTRA if interpretation or further instructions are needed.

## 2.14   Setting check up:

CAUTION!

If the analyser is currently being controlled by the PC, disable the latter by clicking on

the disconnect button of the main window.

- Press                 ,  then

- Ignore the error message

- Press the number of the measuring range to be examined, e.g.

0.000035 24300 0

Now the display shows the setting as above. Starting from the top, the following

information appears:

- The measuring range

- The calibration factor

- The linearisation factor

- The blank value

- Conclude this section, by pressing

CAUTION!

Do not press any other key instead of Enter, otherwise the linearisation factor will

be affected and the analyser will be totally inaccurate.

## 2.15   Thermocouple testing:

Disconnect the plug J1 from the board TH 43 and measure the resistance of the

thermocouple with a digital multimeter, range 100 Ohm.

The resistance must be very low  about 0,5 Ohm.

If the resistance is very high or infinite, then the thermocouple is faulty and must be replaced.

## 2.16   Message “Pump vs maximum”

- Set the main power switch to position 1.

- Set the multimeter to 1000 V DC ( or 750 V DC ), and connect it between the test points TP1 and TP2 ( GND ) of the board NR 32. This is the output voltage to the pump.

CAUTION: The NR 32 is not insulated from the main power. The GND of the NR 32

is connected to the neutral of the main power. It is not the same GND as the GND

of the other boards.

- Set the power switch to position 2. The pump must be running if the oxygen pressure is over 1 bar. A reading of 1.5 bar on the front panel pressure gauge is normal.

- Read the voltage on the multimeter. It is usually between 130 V DC and 160 V DC.

- Squeeze the tube ( 1 ), the voltage of the pump is increasing.

- If not, adjust with P3 of the board FC 21

# 3 SERVICE

## 3.1   IR-paths, cleaning and replacing:

- Unscrew and remove the tubes from the IR-paths.

- Remove the nuts ( 4 ) ( Size: 7 ); in doing this, the threaded rods ( 6 ) should be held

tight to prevent them from turning with the nut.

To be able to remove the nut ( 4 ) from the maximum length IR-path (10),

remove the plate ( 1 ) by unscrewing the four screws ( 2 ).

Take out the pre-amplifiers ( 7 ) together with the spacers ( 8 ).

- Clean the paths with a cleaning solvent, e.g., ethanol.

- With a three part path, consisting of ( 12, 13, 15 ) the windows can be cleaned on the inside using a cotton swab.

- In case of a single part path, such as ( 10 ), the path is filled with ethanol and, after a few minutes, oxygen is blown in, to clean and dry out the path.

Soiling, which can accumulate from the analysis of materials that contains a lot of sulphur and moisture or hydrogen, can be very difficult to clean. In such cases, the paths should be replaced.

- Before reassembling the paths, the threaded rods ( 6 ) should be screwed in by hand, in order to ensure that they are firmly installed.

- Assembly of the paths follows in the reverse order.

- Reconnect the plugs ( 17 ) of the preamplifiers ( 7 ), in case they were removed.

- After reinstallation, adjustments according to 2. 3 are highly recommended.

How CO2 detectors can be recognized:          How SO2 detectors can be recognized:

Square window                                                  Round window

Initial option:                     without any                 Initial option:                   without any

type number.                                                      type number.

Round window                                                   Round window

Later option:                     with type No.:              Later option:                   with type No.:

407M36-611                                                       407M36-118

- In case the infrared detectors ( 18 ) are to be replaced, you can recognize them as described above.

ATTENTION!:

After replacing a specific IR-detector, it is compulsory to zero-adjust:

- see 2. 3 under „ Adjusting the IR cell “

- linearise, see Operation Manual 5. 5 when operated by PC, see 2. 9

and calibrate the corresponding measuring range:

- see Operation Manual 5. 5 when operated by PC see Operation Manual chapter 6.

- black	marked tube connection

- yellow	marked tube connection

- red	marked tube connection

- by reinstalling the infrared cell unit in the analyser, you must pay attention to the colour markings of the tubes and their connections.

Available infrared paths for the CS-2000:

| Path length | % C | Material/ Sample weight | % S | Material/ Sample weight |
| --- | --- | --- | --- | --- |
| 1 mm |  |  |  |  |
| 3 mm | 6 | Metals / 500 mg |  |  |
| mm |  |  |  |  |
| 6 mm | 5 | Metals / 500 mg |  |  |
| 10 mm |  |  | 40 | Non Metals / 200 mg |
| 20 mm |  |  | 20 | Non Metals / 200 mg |
| 30 mm |  |  |  |  |
| 40 mm |  |  |  |  |
| 50 mm |  |  |  |  |
| 100 mm |  |  | 4 | Non Metals / 200 mg |
| 120 mm | 0,2 | Metals / 500 mg |  |  |
| 140 mm |  |  |  |  |
| 170 mm |  |  |  |  |
| 200 mm | 0,1 | Metals / 500 mg |  |  |
| 275 mm |  |  |  |  |
| 300 mm |  |  | 0,3 | Metals / 500 mg |

“Comments”

Induction furnace operation:

1,5 g Tungsten accelerator are needed when analysing Metals.

Above sample weights for Metals and Alloys are just guidelines. For more accurate information please turn to Operation Manual 3. 3 „Applications“

The following materials are used as Non Metals: cement, stone, soils, etc.

## 3.2   IR-source:

- The IR-sources ( 2 ) are fastened

with the screws ( 3 )

- To install or removing the IR-source

( 2 ), the complete chopper

motor assembly ( 1 ), together with

the mounted paths and pre-amplifier,

should be removed from the infrared

unit, see 3. 5

The chopper motor housing ( 1 ) does

not need to be unscrewed. Either by

removing or installing the IR-source.

- In order to remove the IR-source,

unscrew the nuts ( 4 ) with a

( 5.5 mm wrench )

and counter the screws ( 3 ) at the

Same time.

CORRECT					 WRONG!

CAUTION !

The IR source ( 5 ) itself should be in the middle of the ring ( 6 )

( this is without considering the location of the soldered connections,

which are not centered ).

See above drawing: A and B are correct, C, D, E and F are wrong.

Avoid moving the IR-source back and forth, try placing it correctly the first time.

- After re-assembling the infrared cell, adjust according to 2. 3

## 3.3   Infrared electronics:

The complete signal processing for all channels is handled through the circuit board IR 41. In the event of a failure, the entire circuit board should be exchanged.

- The plugs ( 1 / 2 / 3 / 4 / 5 / 6 ) should be unplugged.

- The 10 screws ( 7 and 8 ) should be removed.

- No adjustment is needed on the IR 41-circuit board. You should, however, check the adjustments on the infrared cell, see 2. 3.

The IR 41-circuit board is malfunctioning, when the correct signal can be measured at test point 1; while the signal at test  point 6 cannot be adjusted to zero, see 2. 3

## 3.4   Infrared power supply:

- The power supply ( 1 ) is a complete unit, which is regarded as a single spare part.

- If there is a failure, the entire module ( 1 ) is replaced.

- In order to keep costs down, an exchange of modules is offered.

Replacement:

- Remove the four screws ( 2 ) from the front panel ( 1 ).

- Remove the two screws ( 3 ) from the front panel ( 1 ), to loosen the internal plug ( 4 )

- Unplug the 2nd internal plug ( 5 )

- The front panel is now completely loose and can be exchanged.

- The new power supply unit ( 1 ) is installed in the reverse order.

- It is suggested that the adjustments now be made according to 2. 4.

## 3.5   Chopper:

see  3. 1

The chopper blade ( 1 ) is fastened to the motor ( 2 )

the chopper blade must be removed first  to dismount

the motor ( 2 ) from the mounting plate ( 3 )

In this numeral sequence disassemble the pieces as follows:

- safety ring ( 4 )

the washer(s) ( 5 ) ( there could be two )

the chopper blade ( 1 )

the cylinder ( 6 ) and the

spring ( 7 )

- Reassemble in the reverse order.

CAUTION !

Before re-assembling, check the position of the IR-source ( 8 ) and, if necessary,

centre its position, see 3. 2. The heating wire is not very elastic, therefore do not

move it back and forth, rather, you should correctly position it the first time

- Disassemble the plug ( 1 ) by removing the screws ( 2 ).

- The cables ( 3 ) of the IR source ( 4 ) are then unsoldered from plug ( 1 )

- Unscrew the nuts ( 5 )

- Install the new motor ( 6 )

- The IR source cables ( 3 ) are then soldered to the plug ( 1 )

CAUTION !

Due to the high current for the infrared cell;

it is important that the soldering here is done accurately, because two pins inside the

plug ( 1 ) will be soldered together to form one group. There are two such groups in

one plug.

Each end of the IR-source cables ( 3 ) will then be soldered to each one of these groups. Polarity is irrelevant.

- Afterwards, the chopper blade is refastened.

- Reassemble in the reverse order.

- After installing, the infrared cell should be adjusted, see 2.  3

## 3.6   Infrared temperature regulation:

- The temperature regulation of the infrared cell are made up of;

the circuit boards TH 62 ( 1 ) and NR 32 ( 2 )

the connecting wires ( 3 )

the temperature sensor ( 5 ) and the

heaters ( 4 ).

Replacing the TH 62 :

- The plug ( 4 ) should already have been disconnected during the removal of the infrared cell from the analyser.

- Unscrew and remove the power supply module ( 1 ), see 3. 4

- Disconnect plug ( 2 )

- Remove the four nuts ( 3 ). Re-assemble in reverse order.

- Make adjusts to the unit, according to 2. 5

Replacing the heaters:

- Unsolder the cables ( 2 )

- Cut the silicon seal away from the heaters ( 1 ) and then pull them carefully out

- Insert the new heaters with smearing thermal conducting paste ( in case there is none available, silicon grease can be used ).

- Solder the cables.

- Secure the heaters against falling out with silicon sealant.

Replacing the temperature sensor:

- Unsolder the cables ( 4 )

- Cut the silicon sealant away from the temperature sensor ( 3 )

- Screw the new temperature probes in

- Solder the cables ( 4 )

CAUTION !

screw the probes in very gently, the threads are of aluminium.

- After replacing the temperature sensor, adjust according to 2. 5

## 3.7   Leak checking:

- Set the power switch ( 1 ) to pos. 2

- Close the furnace

- Press and hold the button ( 2 )

The entire system will be checked for leaks.

- The leakage test is completed when  the initial pressure drops (about 5 seconds) on gauge ( 3 ) and then remains constant, the gas system is ok.

- Should there be a continuous pressure drop, then release the button ( 2 ),

open the furnace, press and hold the button ( 2 ) again. If the pressure is

stable then the leakage will found in the furnace.

- If the pressure still decreases, then the leakage is to found inside the analyser.

- Release the button ( 2 ), close the furnace, squeeze the tube ( 4 ) tightly and press and hold the button ( 2 ).

- If the pressure on gauge ( 3 ) remains constant, then the furnace has to be checked for leaks.

- If the pressure shown on gauge ( 3 ) drops, then the leakage has to be found somewhere along the furnace inlet system.

Leaks in the furnace inlet system:

- After following the above instructions, check for gas leakage between inlet tube ( 6 )

and reverse valve ( V2 ).

Leaks in the furnace:

- After following the above instructions, close the furnace, squeeze tube ( 5 ) tightly, press and hold the button ( 2 ), observe the pressure gauge ( 3 ).

- If the pressure drops, then the furnace is leaking. Check whether the O-rings ( 9 ),

( 10 ) and ( 11 ) are soiled or defective see Operation Manual 6.11

Check whether the combustion tube ( 12 ) is broken or cracked.

- If the pressure remains constant, then the furnace outlet system is leaking.

Leaks in the furnace outlet system:

- After following the instructions in the section „Leaks in the furnace“, check if the handle ( 8 ) is properly shut, or else there will be a major gas leak from the dust filter. Check the following components for leakage, by squeezing the following tubes,

while pressing and holding the button ( 2 ):

Tube  to squeeze:	Leak  location:

24		Dust trap ( 6 ), glass tube ( 7 ), dust filter cartridge ( 23 ),

outflow valve ( V 1 )

Leaks inside the analyser:

- when pressing the button ( 43 ), the gauge ( 12 ) shows a pressure drop while the furnace is open, then locate the leakage, according to the sequence below.

Keep the furnace opened, squeeze the following tubes tightly, while pressing and holding the button ( 43 ):

Tube to squeeze:	Leak location:

13		Exhaust valve ( V9 )

14		Flow meter ( 40 )

15		CO2-Paths ( 38 )

16		SO3-Trap ( 35 ) or oxidising furnace tube ( 36 )

17		SO2-Path ( 32 ) or flow sensor ( 33 )

18		Regulator valve ( V6 )

19		Pressure regulator ( PR3 )

20		Bypass valve ( V4 ), check for 24 V DC on input pins

21		Purging valve ( V3 ), check for 0 V DC on input pins

22		Reverse valve ( V2 ), check for 24 V DC on input pins

ATTENTION:  do not cause a short circuit while measuring!

- If the voltages on V2, V3, or V4 are wrong, then replace the electronic drivers, see 1. 6.

## 3.8   Replacing EPROMs:

- Set the power setting to pos. 1

- Open the furnace.

- Pull out all of the plugs from the electronic unit

- Open the left side of the analyser.

- Unscrew the electronic unit.

- Pull out the electronic unit.

- Carefully, take out the two EPROMs H + L with the aid of a screwdriver ( using it as a lever ), from the CPU board. See drawing below.

- Insert the new EPROMs H + L, notice the comparing marks. See drawing below.

Be sure that they are positioned correctly.

Important: When installing

27C512 EPROMs, check that

the jumpers are connected

as shown in the drawings.

If the instrument has an older

card, follow the first drawing;

else follow the second drawing.

- Exchange the RAMs with each

other to ensure that after switching

on the electronic unit, any stored

data during operation with the

previous EPROMs, will be deleted.

CAUTION: The analyzer has to be

recalibrated, or an application has

to be recalled from the application

memory.

## 3.9   Clearing memory:

- Set the power switch to pos. 1

- Open the furnace

- Detach the left hand cover of the Analyser

- Detach the power plug ( 1 ) from the electronic unit

- Press the reset key ( 2 ) on the base plate of the electronic unit  for 30 seconds

- Restore the unit to its original condition

- Recalibrate.

## 3.10   Electronic reset:

Should any problems arise with the electronic unit:

- Press STOP

If the problem remains :

- Turn the power switch to pos. 1

- Open the furnace

- Disconnect the cable ( 1 ) from the electronic unit for 10 seconds and plug back in.

- Should the problem remain :

see 3. 9 " Clearing Memory "

## 3.11   Coding IR-Ranges:

## 3.12   Adding new IR-Ranges:

( from 1 to 2 , from 2 to 3, from 3 to 4 or form 1 to 4 channels )

- The chopper housing with the appropriate number of IR-sources.

Before installing, the IR-source positions should be checked and, if necessary, repositioned, see 3. 2 and 3. 5.

- One, two or three additional pre-amplifiers needed

- One, two or three additional infrared detectors with the appropriate interference filters needed

- One, two or three additional IR-paths of the appropriate length needed

- On the circuit board IR-41, cables must be rearranged correctly. see 3. 3

- Connection of the transformer, matching the number of channels, see 2. 4

" Setting the IR-source Voltage "

- Adjustments for the IR-source power, respective for each channel.

see 2. 4 " Setting the IR-source Voltage "

- Adjusting the zero,

see 2. 3 " Adjusting the Infrared Cell "

- Coding of the channels,

see 3.11 " Coding of the Channels "

- The settings for the temperature regulation,

see 2. 5 " Temperature regulation "

- Afterwards, the linearisation of the Analyser,

see 2. 9 " Linearisation "

- Recommended spare parts. See 4.1

The following components may be needed, when adding more IR cells:

05257		Chopper housing, with one IR-source and motor

05254		Chopper housing, with two IR-sources and motor

05251		Chopper housing, with three IR-sources and motor

05244		Chopper housing, with four IR-sources and motor

05066		IR-path: carbon low

05067		IR-path: carbon high

05068		IR-path: sulphur

06741		Printed circuit board IR-41

06032		Infrared receiver with CO2 filter

06034		Infrared receiver with SO2 filter

## 3.13   Cleaning the solenoid valves:

Valid for  2/2 valve			      	 Valid for 3/2 valve

- Turn the power switch to pos.0

- Unplug the main power plug as well

- Remove the connector ( 7 )

- Unscrew the nut ( 1 )

- Remove the coil ( 2 ) ( only with the 2/2 valve )

- Unscrew counter-clockwise the hexagon screw on the armature housing.

- Remove the armature ( 5 ) ( only with the 2/2 valve )

Careful ! Do not loose the spring ( 4 )

- With oxygen pressure or compressed air, clean the inside of the armature housing

( 3 ), as well as the armature ( 5 ) and re-assemble the solenoid valve.

With the furnace closed and renewed oxygen pressure, check to ensure that the solenoid valve is sealed, see 3. 7

## 3.14   Replacing the flow sensor:

- Set the power switch to pos. 0.

- Remove the left side panel of the analyser.

- Unplug the two cables at the rear-side of the infrared rack.

- Pull out the infrared rack from the analyser.

- Remove the left side cover ( 1 ), the top cover ( 2 ) and the bottom cover ( 3 ) of the infrared rack.

- Remove the tubes ( 4 ), ( 5 ), ( 6 ), ( 7 ), ( 8 ) and ( 9 ) from the flow sensor ( 10 ).

- Remove the two screws ( 11 ) underneath the flow sensor.

- Remove the old flow sensor ( 10 ), install the new one.

- Reassemble the infrared rack in reverse order.

Pay particular attention to the position of the tubes ( 4 ), ( 5 ), ( 6 ), ( 7 ), ( 8 ) and    ( 9 ).

- Adjust according to 2. 1.

## 3.15   Installing an AC power stabiliser:

CAUTION:  The resistance furnace should never be connected to the voltage stabiliser, but directly to the main power.

When the power voltage sinks below 180 V AC ( even for a short time ), a stabiliser is recommended.

The Analyser can draw, during combustion and according to the power setting, up to 13 A.

The resistance furnace ( 14 ) up to 20 A during heating up.

The maximum current depends on the type and size of the sample and accelerator.

To stabilise such a high current is not easy and not necessary.

1	Analyser			  8	Electronic unit

2	Balance			  9	Power plug

3	Printer				10	Triple plug

4	Computer			11	Voltage stabiliser

5	Monitor			12	Original infrared cell power cable

6	Power strip			13	Original electronic unit power cable

7	Infrared cell			14	Resistance furnace

It is sufficient that the electronic unit ( 8 ) and the infrared cell ( 7 ) are protected by

the stabiliser ( 11 ). This  can easily be done, as both units ( 7 / 8 ) use standard plugs.

It is then advisable to protect the balance ( 2 ), the printer ( 3 ) and the computer ( 4 )

with a stabiliser ( 11 ).

- The stabiliser ( 11 ) is plugged into the triple socket ( 10 ) ( rear of the Analyser ).

- The electronic unit ( 8 ) and the infrared cell ( 7 ) are plugged into a power strip ( 6 ), using a normal power cable ( the device cable ).

- The original power plugs ( 12 / 13 ) of the units ( 7 / 8 ) are simply unplugged and left hanging.

A simple transformer can be used when the voltage of the power line is always too low or temporarily, lower than 200 V AC, instead of a stabiliser.

The power of the transformer should be at least 100 VA. The primary voltage should be 220 V to 240 V. The secondary voltage should be about 20 V higher than the primary voltage. For example  Pr 220 V / Sec 240 V or Pr 230 V / Sec 250 V.

The power of the stabilizer should be about 250 VA. The IR-cell and the electronic unit take less than 100 VA, however the power drawn by the PC, monitor, balance and printer, depends on the models connected. Sometimes these devices are provided by the customer himself, so that we can not exactly define the power of the stabilizer. It is advisable to check the power of the above devices before ordering a stabilizer.

## 3.16   Additional safety features for induction furnace pneumatics:

(optional)

The piston doesn’t change position in case of a power failure:

By adding a solenoid valve ( 8 ) and connecting the coil directly to 220 V AC, the piston won’t change position in case of a power failure. If the piston was down, it will remain down; if it was up, it will remain up.

## 3.17   Replacing the gas flow pump:

Replacing the previous model by a new pump of the same model:

- Switch off the instrument and pull the main plug

- Disconnect plug ( 1 ) from the pump ( 2 ).

- Disconnect tubes ( 3 ) and ( 4 ) from the pump

- Remove plastic rings ( 5 ), remove the pump

- Reinstall pump in reverse order

- Replacing by the new model:

The above pump can be replaced by the

later and better model shown below.

Comparing to the previous model, this model is

able to hold the correct flow rate in higher flow

restriction.

The support ( 6 ) of the old pump must be removed

before installing the new pump.

CAUTION:  If the new pump doesn’t  run, exchange the brown and blue wires.

Under no circumstances change the yellow / green wires.

## 3.18   Replacing the heating elements:

- Remove both sides

panels of the furnace

cabinet.

- Remove the cover

plate ( 1 ) of the

cabinet ( 4 screws on the top )

- Disconnect the oxygen

tube ( 3 )

- Unscrew the two knurled nuts ( 4 )

and remove the front assembly ( 5 )

- Remove the O-ring ( 6 ) and the

safety spring ( 7 ) from the front end

of the ceramic tube

- Pull the dust trap ( 8 ) with the

ceramic tube ( 17 ) out of

the furnace

- Remove the spring ( 9 )

and pull the thermocouple ( 10 )

out of the furnace

- Disconnect the

connectors ( 11 ) of the

temperature switch and

pull the cables out of the hole

( 12 )

- Disconnect the cables ( 13 )

- Unscrew and remove the four

screws ( 2 )

- Lift the top frame ( 14 )

together with the furnace

block ( 15 )

- Remove the clamps ( 16 )

from the heating elements

- Install the new heating elements

- Remove the ceramic spacers ( 19 )

- Re-assemble in reverse order

CAUTION

- Install the clamps ( 16 )  in a

position which allows the heating

elements to move at least 5 mm

in horizontal (axial ) direction.

- Insert the thermocouple deep enough,

to touch the combustion tube and fix

it with the spring ( 9 )

## 3.19   Replacing the combustion tube:

CAUTION !  Unplug the main power cable before opening this device !

- Switch off the furnace and disconnect the power plug

- Remove the tube ( 1 )

- Unscrew the screws ( 2 ) and remove the front assembly ( 3 )

- Remove the o-ring ( 4 ) and the safety spring ( 5 )

- Pull the dust box ( 8 ) out, together with the combustion tube ( 6 )

out of the rear of the furnace

- Remove the boat stop ( 7 ) out of the combustion tube

- Dismantle the dust box by removing the screws ( 10 )

and, remove the cover ( 11 ) using a screw driver if

necessary (12 )

- Unscrew the screws ( 13 ) and remove the complete

dust box assembly ( 14 )

- Remove the old combustion tube ( 6 ) install the new one.

CAUTION!

Slide the new combustion tube very slowly and perfectly horizontally

into the furnace or else the heater elements may get damaged.

- Slide the boat stop ( 7 ) into the

combustion tube and ensure that it is

positioned according to this drawing,

( furnace rear-side )

- Reinstall furnace in reverse order.

## 3.20   Installing the halogen trap:

## 3.21   Gas saving mode:

## 3.22   Gas saving also for resistance furnace operation:

If this feature is required, the CS-2000 can be upgraded as follows. But the customer must take care to have oxygen flow first before entering the sample into the furnace. The flow is restored as soon as any button on the PC-keyboard is pressed.

For example taking a sample weight.

If an organic sample is entered into the furnace without flew and the flow is subsequently restored, it may lead to a very intensive ignition. Therefor, the gas saving feature for the resistance furnace operation should be installed only after the customer’s consent to take full responsibility for entering a sample into the furnace only after the flow is restored.

The modifications are shown in the following schematics. For full

Details, see 4. 2.

1.	Install a solenoid valve No. 01160 in the analyser’s cabinet.

2.	Connect the valve in the tube leading to the resistance furnace inlet ( close one port     of the T-piece of the valve, like already done for V3 ).

3.	Connect a cable to pin 49 of the 50-pin plug to the board HF-42.

- Connect the coil of 01160 as shown in schematic.

## 3.23   Upgrade for copper analysis without accelerators:

The CS-2000 ( CS-800 ) can analyse Copper of very high weight, without using any

accelerators. Copper sample up to 20g, or as much Copper sample the crucible can

accommodate.

If the Copper sample has the form of grains or drillings, the standard CS-2000 ( CS-800 )

can do the job.

The condition is:

The furnace control board should have a HF-42 with the corresponding jumper setting for

The correct control mode. See 1. 4

If the Copper sample is a solid piece then the temperature may not go high enough to reach

the melting point. In this case the following modifications are necessary:

WARNING ! High voltage

- Switch off the analyser and unplug the main power cable.

- A second capacitor 500 pF ( ordering no. 77330 ) will be added between the adjusting

frame ( Copper band: ordering no. 13135 ) and the metal mountung frame. The distance of the capacitors to the frame, has to amount to 8 mm min. See drawing below and next page.

- Install a second 100 pF capacitor ( ordering no. 77320 ) parallel to the first

one between the cooling system and the coil, mounted on a U-form metal frame.

( ordering no. 13085 )

- Place a high voltage capacitor ( ordering no. 96V/ 700 10 ) on top of the oscillator’s

housing, fixing it with tape. A red cable ( + ) will be joined from the high voltage

capacitor to the lower capacitor that is fixed on the bottom housing.

A black cable also from the high voltage capacitor ( - ) will be jointed to the capacitor

coil, also mounted on the bottom housing.See drawing below.

## 3.24 Upgrading of analysers with electronic rack for using the new

Windows software:

Late analysers equipped with the new flow sensor (pressure transducer) can easily be upgraded as follows:

( Remove the electronic rack.

( Install the new hardware at the same place.

( Connect the adapters and the new cables for PC and balance.

( Install the front frame and the rear frame.

Upgrading of older analysers equipped with the previous flow sensor (rotameter in the IR-cell) is theoretically possible but practically difficult. The new software is made for the latest versions of analysers with the electronic flow sensor and with two oxygen inlets furnaces. Therefore, older analysers can not run with the new software, unless the analyser is first modified to correspond to the later version. This means replacement of the wiring, new HF board, new upper furnace part with two oxygen inlets and modifications of the tubing and new flow sensor in the IR-cell.

# 4 MISCELLANEOUS

## 4.1   Ordering numbers:

See Operation Manual “chapter 8“

05020		Chopper blade

05030		Chopper blade holder

05065		Chopper motor

05048		Infrared source ( emitter )

70280		O-ring

70330		O-ring

75120		Spring

75130		Retaining washer

75190		Washer

Left side:

Separator plate side

09271		Gas pump

11180		Dust filter cartridge

11370		Gas flow regulating valve

11380		Pneumatic valve

11390		Oxygen solenoid valve

11400		Pressure outlet solenoid valve

11415		Oxygen stop solenoid valve

11415		Lance valve

11430		Purge solenoid valve

11440		Bypass solenoid valve

11490		Pressure regulator

11492		Inlet pressure regulator

11495		Purge pressure regulator

12016		Gas flow and furnace control board HF 42

12044		Transformer

16110		Power supply board NK 31

24250		Valve V10

25190		Valve V11

24260		Valve V13

24265		Valve V12

24285		Board CS20

35330		Attenuator volume

Right side:

Ventilation

11190		Exhaust muffler

11380		Pneumatic valve

11420		Coil cooling solenoid valve

11492		Pressure regulator

12045		Transformer

12080		Rectifier

12100		Transformer

44500		Centrifugal blower

77050		TRIAC

77135		Capacitor 2 µF

Oscillating circuit:

13080		Capacitor support		77140		HF - filter 250V

13090		Upper coil connector	77210		Oscillator tube

13100		Lower coil connector	77320		Capacitor 100 nF

13110		Anode connector		77330		Capacitor 500 pF

13120		Capacitor connector	77335		Capacitor 600 pF

13130		Capacitor connector	77337		10 A Filter

13140		Ground connector		77350		Capacitor 100 nF

13150		Anode heat sink		77600		Resistor 2.7 K(, 100 W

13160		Coil heat sink		77610		Resistor 100 K(, 50 W

13170		Radiation shield

13175		Insulator

13210		Grid choke

13220		Anode choke

13250		Chassis support

13260		High voltage filter

13261		Capacitor 600 pF

13262		Capacitor 600 pF

13270		Resistor 22 (

Induction furnace:

13067		Combustion coil

14009		Pneumatic cylinder for furnace lift

14021		Upper furnace lock

14025		Lower furnace lock

14090		Bearing

14100		Mounting rod

14130		Combustion tube

14161		Lower knurled nut

14165		Upper knurled nut

14168		Pedestal

14170		Pedestal mount

14180		Furnace closure

14185		Tray

14200		Metal tube

14210		Threaded rod

70380		O-ring

70390		O-ring

90150		Crucibles od =25.5; h =27.5 mm

Furnace cleaning:

14021		Upper furnace lock

14045		Cleaning brush for combustion tube

14051		Brush holder

14072		Ceramic heat shield for brush

14080		Cleaning mechanism rod

14161		Lower knurled nut

14165		Upper knurled nut

70120		O-ring

75122		Spring

75130		Safety spring

75150		Metal tube

Oxygen purifying furnace:

21010		Heather section

21120		Quartz tube

70380		O-ring

Front side:

09090		Reagent tube			70210		O-ring

11105		Metal dust filter			70230		O-ring

11062		Reagent tube			70320		O-ring

11064		Reagent tube			70350		O-ring

11480		Adjustable restrictor			70370		O-ring

14013		Induction furnace			72010		Pressure gauge

15083		Gas flow indicator 15 l/h		72020		Pressure gauge

15087		Gas flow indicator 300 l/h		77430		Panel meter 100 µA

15095		Gas flow indicator 600 l/h		78010		Main power switch

18175		Electronic rack			78017		Main power switch

20040		Catalyst tube				78040		Button for leakage test

35355		Lower reagent tube connector

36445		Platform for combustion boat

36700		Upper reagent tube connector

Rear side:

05390		IR-cooling system

11035		Cooling fan

11492		Pressure regulator

36620		Dust system

36730		Viton tube

36735		PVC tube

Pump section:

09271		Gas pump

11370		Flow regulating valve

15037		Pump control card

24260		Pressure switch

35330		Attenuator volume

35331		Attenuator volume

Dust trap cleaning:

36720		Tube

36725		Supplied Teflon tube

for the cleaning of dust deposits

inside the tube 36720

Furnace:

36101		Boat stop

36216		Combustion boat insertion stick

36217		Combustion boat removing stick

36242		Ceramic plate

36283		Heating element connector

36288		Heating element connector

36445		Front panel for the combustion boat

36465		Ceran plate

36620		Dust trap

36665		Spring

36799		Cooling fan

* 36907		Temperature control board TH 43 / TH 44

36908		Cable for TH 43 / TH 44

36914		Thermocouple

36975		LCD – display DVM 2

70380		O-ring

70410		O-ring

75140		Coil spring washer

77450		Temperature switch

77507		Ceramic tube

77525		Heating elements, 4 pcs

90155		Combustion boats

90162		Combustion tube

* Note:

36907   Temperature control board TH 43 is an older version.

36904   Temperature control board TH 44 is a newer version and fully compatible.

Dust trap:

07320		O-ring

07350		O-ring

07370		O-ring

10169		Glass support

11091		Dust trap mechanism

11093		Filter plate

11105		Metal dust filter

11110		Filter housing

36751		Panel

## 4.2   Wiring and gas flow diagrams:

Right  side:

Left side:

Valve wiring:

Control wiring:

IR-current supply:

Attention:

The next four drawings consider

the IR-rack with four IR-cells

The pages after “Connection

between TH 62 and NR 32”

are for the older Model ( S/N 1080

till S/N 1079 )

Infrared cell:

Chopper motor and IR-source:

Connections between TH62 and NR32:

Attention:

Some CS-2000 analysers may be furnished with the previous model of IR-rack, which was designed for maximum three IR-cells. The wiring and the plugs of that rack are different to the rack IR-4 which is described in this manual. The adapter of the drawing below is used to adapt the new IR-rack to the CS-2000 analyser. The drawings of the previous rack wiring can be found in the following 5 pages, these are the older models.

Use the adapter below to install the new IR-cell IR-4 in a previous analyzer wired for the previous IR-cell model.

The EPROMs and the Windows software must also correspond to the new IR-cell. Also the coding of the IR-ranges must match. See 3.11-1.

IR-current supply:

A	13-pin plug connector

B	25-pin plug

C	Transformer

D	Energy filter

E	Heating elements

F	Temperature sensor

*	According to the number of channels, see 2.4

Infrared cell:

1	C-signal

2	GND

3	S-signal

4	Control voltage C

5	Control voltage S

6	Control signal for low channel

7	Regulating voltage for low channel

Chopper motor and IR-source:

Connection between TH 61 and NR 32:

Connecting cables:

Analyzer-Parallel printer

Analyzer-Sartorius balance

Analyzer-PC

CAUTION! Connect the shield on both connector frames.

Gas flow diagram:

Gas flow diagram:

Gas flow diagram:

Analogue I/o signals of the electronic unit:

## 4.3   Starting through external key:

The analysis can be started by means of an external connected key, if desired.

- The key is connected to a 96-pin connector on the rear-side of the electronic unit.

For this purpose, remove the case of the 96-pin plug and connect a wire to the key

( pins 17a to 26a ). ( All carry +24 volts )

- The second wire is connected to pin 5C.

## 4.4   External reports of analyses condition:

The analysis is running; while the word "ANALYSIS" stands on the display; the pin 24B on the 96-pin connector on the rear-side of the electronic is switched to zero.

One could, for example, attach a 24 V relay in order to receive a potential-free

contact message that an analysis is underway.

- The relay is attached to the 96 - pin connector, on the rear-side of

the electronic unit. To do this, remove the cover of the 96-pin

connector and attach a relay wire to 17a to 26a

( all carry +24 volts ).

- The second wire is attached to pin 24B.

## 4.5   Spare parts kits:

For the CS-2000 it is advisable to take both, the CS-500 and the CS-800 kit instead of having a separate kit for the CS-2000

Common spare parts kit for all analysers

92610    Tube of grease						1

11064    Reagent tubes 280mm x 16mm			2

11480    Adjustable flow restrictor				1

15087    Gasflow meter 300 l/h					1

15095    Gasflow meter 600 l/h					1

18175    Electronic unit						1

72010    Pressure gauge 2,5 bar ( 37.5 psi )			1

11035    Cooling fan						1

11492    Inlet pressure regulator					1

11390    Oxygen solenoid valve					1

35310    Pressure switch						1

05067    IR path for high carbon					1

05068    IR path for sulphur					1

05150    IR path connector with window source side		1

05160    IR path connector with window detector side		1

05260    IR path tube (80 to 320 mm)				1

05275    Infrared preamplifier					1

06005    Infrared electronics board IR 21			1

06741    Infrared electronics board IR 41			1

06032    Infrared detector with CO2 filter				1

06034    Infrared detector with SO2 filter				1

06421    Infrared power supply assembly IR13			1

06675    Infrared power supply assembly IR45			1

06058    IR Temperature control board TH61			1

06630    IR Temperature control board TH62			1

15037    IR Temperature regulation board NR32		1

11408    Pressure switch						1

15063    Gasflow sensor 300 l/h					1

05042    Chopper motor						1

05065    Chopper motor						1

05048    Infrared source ( emitter )				1

05060    Reflector							1

75120    Spring							1

75130    Retaining washer						1

75190    Washer							1

18150    Printer interface-cable					1

18160    Balance interface-cable					1

18171    PC-cable							1

73020    Tube transparent						5m

73030    Tube black  id = 4  od = 5				5m

73040    Tube black  id = 6  od = 9				1m

90290    Copper oxide						100g

11062    Reagent tubes  160 x 16mm				2

21120    Quartz tube for purification furnace			1

15083    Gasflow meter 15 l/h					1

72020    Gauge 10 bar/140 psi					1

8010    Mains power switch					1

11380    Pneumatic valve						1

16110    Power supply board NK 32				1

77140    HF-filter 250 V						1

15098    Gas flow control board FC 21				1

71004    PLOTON Windows software ( for OH/ON 900 )	1

71006    PLOTCS Windows software ( for CS 500/800 )	1

70280    O-ring  18 x 2 VT   for chopper motor			2

70330    O-ring  21 x 2 VT   for infrared source			6

70120    O-ring 3.4 x 1.9 VT furnace cleaning rods

( CS 800 ) and for furnace platform		1

70150    O-ring   6 x 2.5 VT					1

70210    O-ring   8 x 3.5 VT for catalyst tube ( CS 800 )and

for reagent tube ( OH/ON 900 )	1

70230    O-ring   9 x 3   VT for reagent tubes and dust filter	1

70320    O-ring  20 x 5   VT for dust filter( CS 800 )and for

furnace glass tube ( CS 500 )	1

70350    O-ring  29 x 5   VT for dust filter ( CS 800 )		1

70370    O ring  34 x 3   VT for dust filter ( CS 800 )		1

70380    O-ring  35 x 5   VT for combustion tube ( CS 800 )

and for furnace internal ( CS 500 )	1

70390    O-ring  38 x 5   VT for furnace ( CS 800 )		1

70285    O-ring  19 x 3   VT for furnace ON/OH 900 ( top )	1

70405    O-ring 47.2x 5.7 VT for furnace closure ON/OH	1

70410    O-ring  48 x 3   VT for furnace internal ( CS 500 )	1

70415    O-ring  62 x 3   VT for furnace ON/OH 900 (centre)	1

70425    O-ring  90 x 2.5 VT for furnace ON/OH ( bottom )	1

## 4.6  Maintenance CS-2000:

Is the furnace closure cone still firm? See 2.8

Check whether the imprint of all the O-rings exceeds 2 mm.

Is the inlet pressure regulator set on 1,5 bar ( 22.5 psi )?

Are the other two pressure regulators set on 0,35 bar ( 5 psi )? See 2.7

Check for leaks. See 3.7

Is the gas flow rate correct? See 2.1

Pressure switch: change over at___________  bar. See 2.2

Does the thermostabilisation work correctly? See 2.5

Check the infrared baseline. See 2.3

Check if all the screws of the generator housing are tightened.

Check whether all four „high voltage“ stickers are still in place.

Check if the 96-pole plug at the back of the CS-2000 is firmly plugged in.

Check if all the warning labels are still attached to the side-panels.

Calibration of all measuring ranges.

# Text boxes

If the analyser has the HF 42 board:

Set the jumper setting to connect A1 with A2 on J3

How to set the jumpers for gas saving mode:

Caution:

This option is functional with EPROMS and PC software versions, later than June

2001.

If the analyser has the HF 41 board:

Set the jumper setting to A1 with A2 on J2

Low  carbon