Hydrogen in titanium

Power mode:					Continuously

Sample drop:					Automatic

Out-gassing (Power / time)			3,3 kW for 60s

Purge time:						15s

Stabiliser time:					30s

Integration delay:					10s

Analysis power / time:				2,8 kW for 45s

Min. analysis time:				50s

Accelerator / Flux:					1g Sn for out-gassing in crucible

Calibration:						Standard titanium 100 ppm.

Inner and outer crucible required

100 – 250 mg sample size.

Hydrogen in zirconium

Power mode:					Continuously

Sample drop:					Automatic

Out-gassing (Power / time)			3,2 kW for 45s

Purge time:						15s

Stabiliser time:					30s

Integration delay:					5s

Analysis power / time:				2,8 kW for 45s

Min. analysis time:				60s

Accelerator / Flux:					1g Sn for out-gassing in crucible

Calibration:						Standard titanium or zirconium.

Inner and outer crucible required

Hydrogen in steel

Power mode:					On-Off-On mode

Sample drop:					Automatic

Out-gassing (Power / time)			4,5 kW for 30 - 45s

Purge time:						10s

Stabiliser time:					30s

Integration delay:					10s

Analysis power / time:				3,3 kW for 30s

Min. analysis time:				50s

Accelerator / Flux:					None

Calibration:						Standard steel 5 ppm H

Free hydrogen in aluminium

Power mode:					On-Off-On mode

Sample drop:					Manually

Out-gassing (Power / time)			4,0 kw for 45s

Purge time:						15s

Stabiliser time:					30s

Integration delay:					2s

Analysis power / time:				0,6 kW for 100s

Min. analysis time:				100s

Accelerator / Flux:					None

Calibration:						Standard steel 1 ppm H

gas dose calibrate.

Outer crucible for max. sample weight 1g.

Total hydrogen in aluminium

Power mode:					On-Off-On Mode

Sample drop:					Manually

Out-gassing (Power / time)			3,5 kW for 45s

Purge time:						15s

Stabiliser time:					30s

Integration delay:					5s

Analysis power / time:				1,7 kW for 40s

Min. analysis time:				50s

Accelerator / Flux:					None

Calibration:						Standard steel 5 ppm H

1g sample size / 20s extra purge for

manual loading

Hydrogen and oxygen in oxygen free copper (OF)

Power mode:					Continuously

Sample drop:					Automatic

Out-gassing (Power / time)			3,8 kW for 45s - 60s

Purge time:						15s

Stabiliser time:					30s

Integration delay:					5s

Analysis power / time:				2,4 kW for 20s

Min. analysis time:				40s

Accelerator / Flux:					None

Calibration:						Standard copper 100 pmm O

Standard steel 1 ppm H

Oxygen in aluminium

Power mode:					On-Off-On Mode

Sample drop:					Automatic

Out-gassing (Power / time)			5,0 kW for 45s

Purge time:						15s

Stabiliser time:					30s

Integration delay:					10s

Analysis power / time:				3,2 kW for 30s

Min. analysis time:				50s

Accelerator / Flux:					1g Sn together with sample

Calibration:						Standard copper 300 ppm

250 mg sample size

Oxygen in copper

Power mode:					On-Off-On mode

Sample drop:					Automatic

Out-gassing (Power / time)			4,0 kW for 30 - 45s

Purge time:						10s

Stabiliser time:					20s

Integration delay:					10s

Analysis power / time:				2,8 kW for 30s

Min. analysis time:				50s

Accelerator / Flux:					None

Calibration:						Standard copper 300 ppm  O