## Applications

Oxygen nitrogen in boron nitride ceramics

Power mode:					On Off On mode

Sample drop:					Manually

Out-gassing (Power / time)			5.0 kW for 60s

Purge time:						15s

Stabiliser time:					30s

Integration delay:					2s

Analysis power / time:				50 kW for 60s

Min. analysis time:					100s

Accelerator / Flux:					None

Calibration:						Cer (IV)-oxide (Ce O2) pure with 18.59 % O

and boron nitride(BN) pure with 56,40 % N.

Oxygen and nitrogen in titanium

Power mode:					On Off On mode

Sample drop:					Automatic

Out-gassing (Power / time)			5,8 kW for 45s

Purge time:						15s

Stabiliser time:					30s

Integration delay:					10s

Analysis power / time:				5,0 kW for 45s

Min. analysis time:					50s(250mg)

Accelerator / Flux:					Sample placed in a 1g Nickel basket. Software

for continuous power between the out-gas and

the analysis is required. Subsequently a blank

analysis, without a sample, is required. A short

interval is needed after each analysis, to allow

the furnace to cool down.

Calibration:						The furnace needs to be cleaned every 5

analyses.

Oxygen in lead and lead alloys

Power mode:					On Off On mode

Sample drop:					Automatic

Out-gassing (Power / time)			3.3 kW for 45s - 60s

Purge time:						10s

Stabiliser time:					20s

Integration delay:					10s

Analysis power / time:				2.7 kW for 30s

Min. analysis time:				50s

Accelerator / Flux:					1000 mg of copper (oxygen free copper)

1000 mg of OF copper needs to be melted

inside the crucible during the out-gas phase,

before the sample drop unit automatically

drops 1000 mg of lead into the crucible.

Calibration:						Copper 10 ppm O.

Furnace needs to cleaned after 25 analysis

Oxygen in steel

Power mode:					On Off On mode

Sample drop:					Automatic

Out-gassing (Power / time)			5.0 kW for 30s - 45s

Purge time:						10s

Stabiliser time:					20s

Integration delay:					10s

Analysis power / time:				4,0 kW for 30s

Min. analysis time:				50s

Accelerator / Flux:					None

Calibration:						Standard steel O+N in similar range.

Oxygen in aluminium

Power mode:					On-Off-On mode

Sample drop:					Automatic

Out-gassing (Power / time)			5,0 kW for 45s

Purge time:						15s

Stabiliser time:					30s

Integration delay:					10s

Analysis power / time:				3,2 kW for 30s

Min. analysis time:				50s

Accelerator / Flux:					1g Sn together with sample

1000 mg of tin need to be melted inside the

crucible during the out-gas phase, before the

sample drop unit automatically drops 1000 mg

of aluminium into the crucible.

Calibration:						Standard copper 300 ppm

250 mg sample size

Oxygen in magnesium

Power mode:					On Off On mode

Sample drop:					Automatic

Out-gassing (Power / time)			4,5 kW for 45s

Purge time:						10s

Stabiliser time:					20s

Integration delay:					10s

Analysis power / time:				2,8 kW for 30s

Min. analysis time:				50s

Accelerator / Flux:					1g Sn

Calibration:						Standard copper.

Furnace cleaning after each analysis

Oxygen in copper

Power mode:					On-Off-On mode

Sample drop:					Automatic

Out-gassing (Power / time)			4,0 kW for 30 - 45s

Purge time:						10s

Stabiliser time:					20s

Integration delay:					10s

Analysis power / time:				2,8 kW for 30s

Min. analysis time:				50s

Accelerator / Flux:					None

Calibration:						Standard copper 300 ppm O

Oxygen in oxygen free copper (OF)

Power mode:					Continuously

Sample drop:					Automatic

Out-gassing (Power / time)			3,8 kW for 45s - 60s

Purge time:						15s

Stabiliser time:					30s

Integration delay:					5s

Analysis power / time:				2,4 kW for 20s

Min. analysis time:				40s

Accelerator / Flux:					None

Calibration:						Standard copper 100 pmm O

Standard steel 1 ppm H

Hydrogen in titanium

Power mode:					Continuously

Sample drop:					Automatic

Out-gassing (Power / time)			3,3 kW for 60s

Purge time:						15s

Stabiliser time:					30s

Integration delay:					10s

Analysis power / time:				2,8 kW for 45s

Min. analysis time:				50s

Accelerator / Flux:					1g Sn for out-gassing in crucible

Calibration:						Standard titanium 100 ppm.

Inner and outer crucible required

100 – 250 mg sample size.

Hydrogen in zirconium

Power mode:					Continuously

Sample drop:					Automatic

Out-gassing (Power / time)			3,2 kW for 45s

Purge time:						15s

Stabiliser time:					30s

Integration delay:					5s

Analysis power / time:				2,8 kW for 45s

Min. analysis time:				60s

Accelerator / Flux:					1g Sn for out-gassing in crucible

Calibration:						Standard titanium or zirconium.

Inner and outer crucible required

Hydrogen in steel

Power mode:					On-Off-On mode

Sample drop:					Automatic

Out-gassing (Power / time)			4,5 kW for 30 - 45s

Purge time:						10s

Stabiliser time:					30s

Integration delay:					10s

Analysis power / time:				3,3 kW for 30s

Min. analysis time:				50s

Accelerator / Flux:					None

Calibration:						Standard steel 5 ppm H

Free hydrogen in aluminium

Power mode:					On-Off-On mode

Sample drop:					Manually

Out-gassing (Power / time)			4,0 kw for 45s

Purge time:						15s

Stabiliser time:					30s

Integration delay:					2s

Analysis power / time:				0,6 kW for 100s

Min. analysis time:				100s

Accelerator / Flux:					None

Calibration:						Standard steel 1 ppm H

gas dose calibrate.

Outer crucible for max. sample weight 1g.

Total hydrogen in aluminium

Power mode:					On-Off-On Mode

Sample drop:					Manually

Out-gassing (Power / time)			3,5 kW for 45s

Purge time:						15s

Stabiliser time:					30s

Integration delay:					5s

Analysis power / time:				1,7 kW for 40s

Min. analysis time:				50s

Accelerator / Flux:					None

Calibration:						Standard steel 5 ppm H

1g sample size / 20s extra purge for

manual loading

Hydrogen and oxygen in oxygen free copper (OF)

Power mode:					Continuously

Sample drop:					Automatic

Out-gassing (Power / time)			3,8 kW for 45s - 60s

Purge time:						15s

Stabiliser time:					30s

Integration delay:					5s

Analysis power / time:				2,4 kW for 20s

Min. analysis time:				40s

Accelerator / Flux:					None

Calibration:						Standard copper 100 pmm O

Standard steel 1 ppm H